Vibration isolation in neonatal transport by using a quasi-zero-stiffness isolator

Jiaxi Zhou1, Kai Wang1, Daolin Xu 1, 2, Huajiang Ouyang 3, Yimei Fu 4

1 College of Mechanical and Vehicle Engineering, Hunan University, Changsha 410082, PR China,

2 State Key Laboratory of Advanced Design and Manufacturing for Vehicle Body, Changsha 410082, PR China,

3 School of Engineering, University of Liverpool, Liverpool L69 3GH, UK

4 Department of Obstetrics & Gynecology, Xiangya Hospital, Central South University, Changsha 410008, PR China

Abstract

New-born infants are likely to suffer excessive vibration during neonatal transport, but there is no effective method to attenuate the transmission of vibration from an ambulance floor to an infant. This would be particularly detrimental to premature or sick new-born babies. In this paper, a quasi-zero-stiffness (QZS) vibration isolation method is proposed to improve the comfort of the infant in an incubator. The infant compartment is supported by QZS isolators, which are realized each by combining a pair of mutually repelling permanent magnets in parallel connection with a coil spring, and their parameters are obtained by design optimization to maximize the displacement range with smaller stiffness than that of the coil spring. A lumped-mass model of the transport incubator is developed, and the vibration isolation performance is estimated in terms of displacement and acceleration transmissibility. Numerical results reveal that a comparatively heavy damping is needed to completely avoid the jump phenomenon inherent in polynomial nonlinearity and suppress resonance, and thus achieve smooth and effective vibration attenuation starting from an ultra-low frequency. Under a random disturbance, magnification of vibration is observed in the original transport incubator, but notable attenuations of Root Mean Square (RMS) acceleration and displacement in the modified one, which indicates that the proposed QZS isolator should be a good solution to eliminate vibration-induced injuries in neonatal transport.

Keywords: vibration isolation, quasi-zero-stiffness, neonatal transport, infant incubator, transmissibility.

- Introduction

Generally, premature or sick new-born infants are transported to perinatal centres for intensive care via ambulances, fixed-wing aircraft or helicopters. No matter what modes of transportation are adopted, mechanical vibration transmitted from the vehicle body to an infant is unavoidable. It has been known from experiments on adults and animals that vibration usually has adverse effects (ISO-2631-1:1997, 1997) on cardiorespiratory functions (Clark et al., 1967) and peripheral and central nerve systems (Floyd et al., 1973; Landstrom and Lofstedt, 1987). Moreover, it was reported that premature infants with very low birth weight were more likely to incur intraventricular hemorrhage and cerebral palsy due to vibratory stresses during neonatal transport (Prehn et al., 2015; Shah et al., 2008). Therefore, mitigating mechanical vibration (and noise) is essential to provide a comfortable and safe transport that prevents premature or sick new-born from avoidable morbidity.

In the past two decades, several attempts (Browning et al., 2008; Gajendragadkar et al., 2000; Prehn et al., 2015; Sallee et al., 2016; Shah et al., 2008; Sherwood et al., 1994) of structural modifications of transport incubators have been made to reduce vibration levels of neonates in the transport incubator. Sherwood et al. (1994) tried to use three kinds of mattresses, namely foam, air-filled and gel-filled mattresses, to decrease vibration intensity. It was found that gel-filled mattress was most effective, and the foam restraint placed between the tray and the track played a vital role to suppress vibration. A similar experiment was carried out by Gajendragadkar et al. (2000), where a combination of gel-filled mattress and foam was considered. The gel-filled mattress was reconfirmed most effective again in reducing vibration levels, but none of these mattresses were capable of attenuating the transmission of disturbances from the ambulance floor to the infant, and the foam mattress even accentuated Root Mean Square (RMS) acceleration by about 3 times during city runs. In contrast, experiment tests carried out by Prehn et al. (2015) indicated that an air-filled mattress was superior to the gel-filled one, which was a different finding from previous works (Gajendragadkar et al., 2000; Sherwood et al., 1994). This can be attributed to the fact that there is no standard for experimental tests and vibration measurements during neonatal transport. In other words, the categories of ambulance, types of transport incubator, and brands of mattress utilized in experimental investigations were different from each other. Browning et al. (2008) found that constraints on the tray had a significant impact on vibration levels during in-hospital transport, which indicated that loosening constraints leads to large vibration amplitude. Furthermore, Salle et al. (2016) fixed the bottom tray onto a rigid frame by bolts to tune its natural frequencies far away from the driving frequency, and inserted rubber strips to dampen resonance.

The above works made an important step towards understanding the predominant frequency components of disturbances from a vehicle body, and the contribution factors toward increase in vibration level. However, due to a lack of theoretical modelling and analysis, those modifications on the incubator are unable to substantially mitigate vibration transmission, even though the vibration level can be reduced by shifting natural frequencies of the transport incubator away from the driving frequencies, which are mostly low frequencies, i.e., natural frequencies of the vehicle (Gajendragadkar et al., 2000; Shenai et al., 1981). Therefore, there is a need for the study and design of more effective vibration isolation devices to provide safer neonatal transport (Gajendragadkar et al., 2000).

To the authors’ knowledge, designs of vibration isolation devices for transport incubators are rarely reported. Beale (2006) designed an active vibration control system to reduce acceleration encountered by infants in incubators. Bailey-Van Kuren and Shukla (2005, 2007) also proposed an active pneumatic-springs vibration isolation system. In both the devices above, supporting instruments and a power-supply are needed to execute the control algorithm. Passive vibration isolators are believed to be more compact, cheaper and more reliable than active ones. Sabota and Aghili (2014) modified an incubator by installing a passive vibration isolation system consisting of coil springs and dampers between the tray and the rigid frame; however, the vibration isolation performance of this system was not estimated in the patent application.

The motivation of this paper is to propose a promising solution to attenuate vibration in neonatal transport by using quasi-zero-stiffness (QZS) isolators. The QZS isolator has intrinsic advantages of isolating low-frequency disturbances over existing passive and active isolators, due to its characteristics of high-static-low-dynamic stiffness and non-power-supply. The QZS isolator can be realized by combining a negative stiffness (NS) mechanism or structure with a positive stiffness element. Several investigations (Carrella et al., 2008; Le and Ahn, 2011; Platus, 1991; Shaw et al., 2013; Sun et al., 2014; Wu et al., 2014; Xu et al., 2013; Zheng et al., 2016; Zhou et al., 2015) have been carried out for compact and reliable NS mechanisms and structures. In this paper, a pair of permanent magnets with mutually repelling feature is utilized to construct a QZS isolator with a coil spring, because of its easily tailored stiffness characteristic by parametric designs and compact structural configuration (Wu et al., 2014; Zheng et al., 2016). A damper is also installed in parallel with the coil spring to suppress resonance and avoid the jump phenomenon. The infant compartment (or housing unit) is supported by QZS isolators to mitigate the transmission of disturbances from both the ambulance floor and auxiliary equipment (or supporting instrumentations) of the incubator. Considering the soft mattress and rubber wheels, the lumped-mass model of the transport incubator is established, and the vibration isolation performance is evaluated in terms of displacement and acceleration transmissibility under harmonic disturbances and a random excitation used in the experimental test by Shenai et al. (1981).

This paper is organized as follows. The lumped-mass model of the original transport incubator is established in Section 2. Section 3 presents the design concept of the QZS vibration isolation system, including design of the modified transport incubator, detailed parametric design and stiffness analysis of the QZS isolator. In Section 4, the lumped-mass model of the modified transport incubator is established, and the vibration isolation performance of the QZS transport incubator system is estimated, which is also compared with that of the original one to demonstrate its effectiveness in isolating low-frequency disturbances.

- Lumped-mass modelling of a transport incubator

## 2.1 Parameters of Lumped-mass model

The transport incubator is equipped in a neonatal intensive care ambulance, consisting of an infant compartment (or housing unit), auxiliary equipment (or supporting instrumentation) and a stretcher (or trolley). The mattress is supported by a tray. The tray is rigidly fixed on the frame that is used to hold the auxiliary equipment. Considering rubber wheels supporting the stretcher and soft mattresses housing the infant, the transport incubator can be approximately modelled as a 2-DOF lumped-mass system excited by disturbances from the ambulance floor, as shown in Fig. 1. Note that only the vertical disturbance is considered here, since this is the dominated motion of the vehicle body of the ambulance, and thus only the vertical DOF of each lumped mass is taken into account.

Fig. 1 Lumped-mass model of the infant-incubator system mounted on the ambulance floor.

The stiffness of a rubber wheel is assessed by using Finite Element (FE) analysis.  Fig. 2a shows stresses distributions of a deformed wheel, which consists of a steel hub with radius of 30.5 mm and solid rubber tire with thickness of 32 mm. The constitutive relationship of rubber is represented by an incompressible Mooney-Rivlin model with parameters C10=3.2 MPa and C01=0.8 MPa, which can be found in ABAQUS user’s manual (2014).  By applying an additional vertical concentrated force to a payload (a quarter of total weight of transport incubator with auxiliary equipment and stretcher) at the centre of the wheel, a further deflection after static deformation can be obtained, and the relationship between the additional force and deflection is depicted in Fig. 2b.  It can be seen that the stiffness is nonlinear, but the nonlinearity is insignificant. Hence, the stiffness of rubber wheel can be obtained from the secant slope of the force-deflection curve, as listed in Table 1.

Fig. 2 Stiffness of rubber wheel. (a) Deformation and stress distributions; (b) relationship between load and deformation.

Table 1 Parameters of the lumped-mass model of a transport incubator

| Parameter | Definition | Value |
| --- | --- | --- |
| (kg) | Mass of the infant | 2 |
| (N/m) | Stiffness of the mattress |  |
| (N/ms) | Damping coefficient of the mattress | 50.2656 |
| (kg) | Mass of the transport incubator | 123.9 |
| (N/m) | Stiffness of the rubber wheels |  |
| (N/ms) | Damping coefficient of the rubber wheels |  |

There are many kinds of mattresses used to house an infant in an incubator, such as gel-filled, air-filled and foam-filled ones. It is difficult to determine the stiffness of these mattresses by FE simulations, due to a lack of knowledge of the specific material parameters. However, it can be approximately identified by using existing experimental data. Gajendragadkar et al. (2000) measured the natural frequency of an infant-mattress system, which was about 10 Hz for a gel-filled mattress, and then the stiffness of this gel-filled mattress can be determined, as listed in Table 1. Note that the mass of the transport incubator in Table 1 is that of Air-Shields TI-500 Infant Incubator (Draeger, 2016), including the mass of a stretcher equipped for ambulance. Damping ratios of gel-filled mattresses and rubber wheels are selected as 0.2 (silicone gel in a solid state (Osada and Kajiwar, 2001)) and 0.08 (polyurethane rubber under average room temperature (Chen et al., 2011)), respectively, and their corresponding damping coefficients are listed in Table 1. It is worth noting that the amplification factors of amplitude at resonance in Refs (Gajendragadkar et al., 2000; Sherwood et al., 1994) are taken into account when the damping ratios are determined.

## 2.2 Displacement  transmissibility

When the ambulance floor is excited by a harmonic disturbance, the displacement transmissibility of the system, namely the ratio of the displacement amplitude of the infant to that of disturbance, is depicted in Fig. 3. Note that the acceleration transmissibility is identical to the displacement transmissibility under a single harmonic excitation in a linear vibration isolation system. As seen from Fig. 3, the lumped-mass model of the transport incubator system has two natural frequencies, namely 10 Hz and 35.21 Hz. Gajendragadkar et al. (2000) reported that the dominant frequencies of disturbance were 2.5 Hz and 15 Hz for a neonatal intensive care ambulance, which were believed to be natural frequencies of the vehicle body. At these two excitation frequencies, the displacement transmissibility in Fig. 3 is 1.071 and 1.025, respectively, which implies no mitigation of transmission of disturbance. Therefore, an effective vibration isolation device is needed to protect the infant from low-frequency disturbances originating from the ambulance body.

Fig. 3 Displacement transmissibility of the transport incubator system.

- Scheme of a QZS vibration system

As mentioned above, the disturbance mainly occurs from vibration of low-frequency components.  Thus, if traditional linear isolators, such as coil springs and rubber mounts, are adopted, their stiffness should be designated to be very low, because this kind of isolator is effective only when its natural frequency is well below the disturbance frequency. Very low stiffness leads to a very large static deflection under payload, which would cause the transport incubator wobbly and unstable. In contrast, quasi-zero-stiffness (QZS) isolators possess the property of high static stiffness but low dynamic stiffness, which not only can support a payload without excessive static deflection, but also have very low stiffness around the static equilibrium position, even close to zero stiffness at equilibrium. The low dynamic stiffness of a QZS isolator is able to attenuate low-frequency disturbances effectively

## 3.1 Prototype design of the QZS transport incubator system

The schematic diagram of the QZS transport incubator system is shown in Fig. 4a.  The compartment (a2) housing the infant is supported by four QZS isolators (a1). Each isolator is rigidly fastened on a frame (a3), which is fixed on the stretcher (a4). The assembly of QZS isolators in the frame is presented in Fig. 4b. Note that the frame with high stiffness makes space for placing auxiliary equipment, such as control devices, flowmeter, gas tank, compressor, and battery. As a result, this QZS vibration isolation system can also mitigate disturbances arising from the auxiliary equipment, especially the compressor.

The QZS isolator is fulfilled by combining a pair of mutually repelling ring permanent magnets in parallel connection with a coil spring, as shown in Fig. 4c. Each corner of the infant compartment is connected to a rigid rod (c1) by a support (b1). The rod (c1) is guided by two liner bearings (c2), and thus it can only move along the axial direction. Inner ring magnet (c4) is fixed on the rod (c1) by fastening screw nuts (c3) at both ends. Outer ring magnet (c5) is fixed on the sleeve (c7) by fastening bolts (c6). Note that these two magnets have the same magnetization direction along the axis of the rod (c1) to realize mutually repelling configuration. The sleeve (c7) is fastened on the frame by bolts and nuts. One end of the coil spring (c9) is screwed on the bottom end of the rod (c1), and the other end is supported on a pre-compression adjustor (c11), which is screwed on the sleeve (c8). To dampen possible resonances and suppress potential appearance of the jump phenomenon of this nonlinear vibration isolation system, a viscous damper (c10) is set in parallel with and inside the coil spring (c9). Under a payload, the inner ring magnet (c4) aligns in parallel with the outer one (c5), while a marked line on the rod (c1) aligns with the top edge of the upper linear bearing. When the magnitude of the payload alters, the pre-compression of the coil spring can be regulated by turning the adjustor (c11) to guarantee that the marked line always stay on the pre-set position, namely aligning with the top edge of the upper linear bearing.

Fig. 4 Schematic diagram of the QZS vibration isolation system. (a) Overview; (b) inside view of the assembly of QZS isolators in the frame housing auxiliary equipment; (c) inside view of the QZS isolator.

By static analysis of the isolator, the restoring force of the QZS isolator as a function of axial displacement x of the rod (c1) can be given by

where the expression of magnetic force was given by Zhong (1998) , and

and

is magnetic pole surface density, and  is the magnetic polarization vector, and  is the unit normal vector;

, the permeability of the vacuum;

, average radius of the inner and outer permanent magnets;

R, the inner radius of the inner permanent magnet;

, width of both the inner and outer permanent magnets;

, height of both the inner and outer permanent magnets;

, air gap between the inner and outer permanent magnets.

The stiffness of the QZS isolator can be obtained by differentiating the expression of restoring force with respect to ,

where

## 3.2 Parametric design of the QZS isolator

Substitute x=0 into Eq.  and let, and then one can derive the zero-stiffness condition at the equilibrium:

By substituting the above equation into Eq. , the restoring force of the QZS isolator can be given by

and the stiffness of the QZS isolator can be given by

As seen from Eqs.  and , the stiffness of the QZS isolator is depended on the stiffness of coil spring and geometric parameters of magnets. To achieve best low-frequency vibration isolation performance, the stiffness is expected to be as low as possible under several constraints, such as restrictive physical dimensions, capability of supporting payload, and zero-stiffness condition. An index  denoting the displacement range with smaller stiffness than that of the coil spring is adopted as the objective function. Through maximizing this index, the optimal geometric parameters of magnets will be found. The mathematical formula for design optimization can be written as

where subscripts U and L denote upper and lower bounds of variables, respectively, and index  is yielded from , namely the positive root of the following equation

and the stiffness of the coil spring is priorly determined according to the level of payload, namely , where is the maximum allowable static deflection of the isolator, and is the payload. Note that under the payload, the isolator is in the static equilibrium state, while the mutual repelling force between the pair of magnets is zero, and thus the payload is supported only by the coil spring.

By installing QZS isolators, the incubator can be divided conceptually into two parts, namely infant compartment and the frame for auxiliary equipment.  The weight of the infant compartment made of polymethyl methacrylate is about, and the total weight of the auxiliary equipment (including stretcher) is about. The maximum allowable static deflection is assume to be 10 mm, and the payload supported by each isolator is 50N, and then the stiffness of the coil spring can be tentatively determined as . Considering space compactness of the isolator, the upper bound of the geometric variables is selected as , and the lower one is select as small positive values.  By carrying out the design optimization by using MATLAB® function fmincon based on the interior-point algorithm, the optimal geometric parameters of magnets can be determined, which is . Taking into account manufacturing feasibility and compliance with the constraint on maximum allowable static deflection, the non-integer value of  is rounded up towards the nearest integer. Therefore, the definitive design parameters of magnets are. Based on these parameters, the stiffness of the coil spring can be determined by Eq. . Finally, parameters of the QZS isolator are summarized in Table 2.

Table 2 Parameters of the QZS isolator

| Parameters | (T) |  | R (mm) | (mm) | (mm) | (mm) | (N/m) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Values | 1 |  | 20 | 20 | 5 | 20 |  |

## 3.3 Stiffness characteristics of the QZS isolator

The restoring force and stiffness of the QZS isolator are shown by solid lines in Fig. 5. As expected, the stiffness is zero at the static equilibrium position, and very low in the vicinity of this position, and below the stiffness of coil spring in a large displacement range. But it has the same capability of supporting the payload like the coil spring. In other words, the payload results in identical static deflections when it is supported by the QZS isolator or only by the coil spring.

Fig. 5 Static characteristics of the QZS isolator. (a) Force and (b) stiffness vs displacement.

As seen from Eq. , the expression of restoring force is so complicated that it is difficult to conduct consequent dynamic analysis. Generally, the original expression is approximated by using a Taylor Expansion near the static equilibrium position. However, this approximation matches well with the original one only in the vicinity of the static equilibrium position, but, outside this displacement range, the difference become so large that the Taylor Expansion cannot describe the stiffness characteristic with acceptable accuracy. Therefore, the tedious expression of restoring force of the QZS isolator is approximated by Polynomial Fittings with components of , and , due to the fact that the restoring force  is an odd function of displacement and the stiffness is equal to zero at .  The approximate expression of the restoring force can be given by

where , , and , which are obtained by using MATLAB® function lsqcurvefit based on the trust-region-reflective algorithm. Both the approximate restoring force and stiffness, which is obtained by differentiating Eq.  with respect to x, are depicted as hollow dotted lines in Fig. 5. It can be seen that there exists excellent agreement between the exact restoring force and its approximation in the whole displacement range from -30 mm to 30 mm, although the agreement of stiffness degrades when the displacement amplitude exceeds 27 mm.

- Dynamic model and simulations of the QZS transport incubator

After four QZS isolators are installed beneath the infant compartment, the incubator is divided into two parts. And thus, the QZS transport incubator system can be represented by a 3-DOF lumped-mass system, as shown in Fig. 6a. As mentioned before, only the vertical DOF of each lumped mass is considered, and the infant compartment can only move along the vertical direction; therefore, all these isolators are assumed to undergo the same amount of deformation. The equations of motion of the 3-DOF model are established, and its vibration isolation performance is estimated in terms of displacement transmissibility, which shows that the wheels have negative effects on the vibration isolation performance. Therefore, the stretcher is suggested to be rigidly fixed on the ambulance floor during the neonatal transport, leading to a modified 2-DOF model, as shown in Fig. 6b. Furthermore, based on this 2-DOF model, the critical damping for totally avoid jump phenomenon and suppressing resonance is determined, and the effectiveness of the QZS isolation system in the modified transport incubator is verified by numerical simulations under harmonic and random disturbances.

Fig. 6 Lumped-mass model of the modified transport incubator. (a) 3-DOF model and (b) modified 2-DOF model.

## 4.1 Dynamics of the 3-DOF QZS system

The equations of motion of the 3-DOF system are given by

where  is the damping coefficient of the QZS vibration isolation stage, and  is a disturbance applied on the ambulance floor. Generally, under a periodic excitation, the QZS system may experience complicated responses, such as sub/super-harmonic, quasi-periodic and even chaotic responses (Zhou et al., 2015). As well-known, under a multiple-harmonic disturbance, the response of a linear system is not periodic with only one frequency component either. For those responses, the definition of displacement transmissibility in Section 2.2, namely the ratio of displacement amplitudes, is unusable. A ratio of RMS displacement of the infant to that of the disturbance is newly defined as displacement transmissibility, i.e.

where   and  are  time histories of displacement responses of the infant and the disturbance, respectively. Time histories of  is obtained by solving Eq.  numerically with Runge-Kutta algorithm.  When a cosine disturbance  is applied on the ambulance floor, both the forward and backward frequency sweeps as linear functions of time are carried out, and the displacement transmissibility is shown in Fig. 7 under different levels of damping and excitation. In this figure, solid lines and dashed lines denote results obtained from the forward and backward frequency sweeps, respectively, and arrows are utilized to distinguish these two types of frequency sweeps from each other when jump phenomenon occurs.

Fig. 7 Displacement transmissibility of the QZS system under different levels of damping and excitations.

From Fig. 7, it can be seen that under low-level excitation the QZS isolation system has a notable effectiveness in the whole frequency range (Fig. 7a), even though an obvious resonant peak occurs at the natural frequency of the infant-mattress subsystem. However, under high-level excitation, the jump phenomenon, a multi-valued feature of nonlinear dynamic systems, is activated, and the amplitude on the upper branch is very large (Fig. 7b). In the vicinity of the natural frequency of the auxiliary equipment-wheels subsystem, the displacement response of the infant is chaotic with large amplitude, as shown in Fig. 8. The frequency spectrum (Fig. 8b) shows that numerous components with frequencies below the excitation frequency  occur, and amplitudes at natural frequencies of two aforementioned subsystems are considerable.

Fig. 8 Displacement response of the infant when . (a) Time history; (b) frequency spectrum.

Moreover, increasing damping ratio from 0.1 to 0.2 is able to attenuate resonance and avoid the jump phenomenon in low-frequency range, as demonstrated in Fig. 7c; however, complicated responses at frequencies close to the natural frequencies of the auxiliary equipment-wheels subsystem cannot be eliminated, which are still in chaotic state. Furthermore, heavier damping is utilized to suppress those complicated responses, as shown in Fig. 7d. Under such a circumstance, the resonance and jump phenomenon of the QZS subsystem are also completely supressed. However, the heavy damping degrades the effectiveness of vibration isolation, although the system can perform vibration isolation in the whole frequency range, as demonstrated in Fig. 7d. To essentially get rid of negative effects arising from the auxiliary equipment-wheels subsystem, the stretcher is suggested to be mounted rigidly on the ambulance floor by lockers during the neonatal transport in the ambulance.  In the next subsection, the vibration isolation performance of the modified 2-DOF QZS vibration isolation system will be studied in details.

## 4.2 Dynamics of the modified 2-DOF QZS system

After the removal of rubber wheels, the lumped-mass model of the transport incubator is reduced to a 2-DOF QZS system, as shown in Fig. 6b. In such a situation, the displacement of the auxiliary equipment (including stretcher) is equal to that of the ambulance floor. The equations of motion of the 2-DOF QZS system can be given by

where   is the disturbance applied on the ambulance floor.

### 4.2.1 Critical damping ratios for avoiding jump phenomenon and suppressing resonance under single harmonic disturbance

Under a harmonic disturbance, by introducing following non-dimensional terms,

the equation of motion can be rewritten in the non-dimensional form

where  is the non-dimensional relative displacement between the infant and his housing compartment, and  is the non-dimensional relative displacement between the infant compartment and the ambulance floor,  and  represents differentiation with respect to . The fundamental response is obtained by using Harmonic Balance Method, and only the primary resonance (the first approximation) is considered.  Therefore, the fundamental response is assumed to be

By substituting Eq.  into Eq.  , and ignoring the higher-order harmonic terms, and equating coefficient of  and , the relationship equations between the amplitudes of fundamental response and excitation frequency can be obtained as

By solving the above equations, the relative amplitudes can be obtained. And then the absolutely displacement of the infant can be given by

where  and  . Therefore, the displacement transmissibility can be yielded,

Variations of peak displacement transmissibility  and the beginning frequency of vibration isolation  against damping ratio are depicted in Fig. 9, by numerically solving Eq. . It can be seen that both  and  decrease as the damping ratio increases in the range of . In addition,  and  are close to 1 and 0, respectively, when the damping ratio approaches a critical value of 0.2675, as depicted by squares in Fig. 9. It indicates that the QZS vibration isolation system will be effective in the whole frequency range, as long as the damping ratio exceeds this critical value.

Fig. 9 Variations of (a) peak displacement transmissibility  and (b) beginning frequency of vibration isolation against damping ratio , when Y=10 mm.

Two key damping ratio values, which are represented by a circle and a square in Fig. 9, respectively, divide the whole range into three characteristic regions. In the range of , obviously, there exists the jump phenomenon, as show in Fig. 10a. When , the jump phenomenon disappears, while a peak displacement transmissibility occurs at low frequency, as shown in Fig. 10b. Therefore, 0.1615 is a critical value of damping ratio for avoiding the jump phenomenon under a harmonic excitation with amplitude of 10 mm. In the range of , the resonance is totally suppressed, and the displacement transmissibility is below 1 in the whole range of frequency, which means that the QZS isolation system is effective under a harmonic excitation with any frequency, as shown in Fig. 10c.

This finding provides a guideline to design a damper to fulfil low-frequency vibration isolation in neonatal transport. It reveals that a comparatively heavy damping is needed to fully avoid the jump phenomenon and suppress resonance, and thus keep the QZS system effective in ultra-low frequency range. Moreover, the comparison between Fig. 10b and Fig. 10c indicates that increasing damping will degrade vibration isolation efficiency when the excitation frequency is out of the ultra-low frequency range.  Therefore, the damping should be selected carefully by considering suppression of resonance and vibration isolation efficiency simultaneously.

Fig. 10 Displacement transmissibility of the 2-DOF QZS system under different damping levels, when Y=10 mm.

### 4.2.2 Vibration-isolation performance under a random disturbance

In this section, the vibration-isolation performance is estimated under a random disturbance

where  is power spectrum density (PSD) of acceleration at frequency , derived from the experimental results during neonatal transports obtained by Shenai et al. (1981), and  is randomly selected based on a uniform distribution in the range from  to , and the value range of  is from 3Hz to 20Hz, and . Unfortunately, in (Shenai et al.,1981), specific values of the PSD were not given, and thus the maximum value of PSD is selected as 0.2 in this paper. To isolate all the frequency components of this random disturbance, the beginning frequency of vibration isolation should be smaller than 3Hz. According to the guide for selecting damping ratio shown in Fig. 9b, the damping ratio should be larger than 0.1675 and it is selected as 0.17 here.

The RMS acceleration is always used to measure vibration levels of human body, and International Standard Organization (ISO) (ISO-2631-1:1997, 1997) evaluates perceptions under different vibration levels based on this measurement. Therefore, the acceleration transmissibility is also an essential index to evaluate the performance of QZS vibration isolation system in neonatal transport, which can be defined by

where  and  are time histories of accelerations of the infant and ambulance floor, respectively. It should be noted that, under a single harmonic disturbance, the acceleration transmissibility is almost equal to that of displacement. But, there exists a notable difference between these two kinds of transmissibility under multi-harmonic disturbances, which depends on the predominated frequency components.

Responses of the original transport incubator, 2-DOF lumped-mass model in Fig. 1, are presented in Fig. 11, where solid lines and dotted lines denote responses of the infant and the disturbance, respectively. It is obvious that amplitudes of displacement and acceleration responses are larger than those of the disturbance. The displacement and acceleration transmissibility, calculated by Eq.  and Eq. , are 1.4580 and 1.8622, respectively, which implies that the transmission of disturbance is amplified by the soft mattress and rubber wheels.

Fig. 11 Responses of the original transport incubator (2-DOF lumped-mass model in Fig. 1). (a) Displacements and (b) accelerations of infant and ambulance floor.

Responses of the modified transport incubator, the modified 2-DOF model in Fig. 6b, are depicted in Fig. 12. In this case, the displacement and acceleration transmissibility are 0.4470 and 0.2775, respectively, which implies that RMS displacement and RMS acceleration transmitted to the infant have been isolated by more than 50 percent and 70 percent, respectively.  In addition, the effectiveness in mitigating transmission of acceleration disturbance is superior to that of displacement, due to the fact that the displacement and acceleration of a disturbance are predominated by lower-frequency components and higher ones, respectively. Form the standpoint of promoting amenities (measured by RMS acceleration) of the infant during neonatal transport, the QZS isolation system provides a feasible solution to isolate disturbances in terms of RMS acceleration, although the effectiveness in attenuating transmission of displacement disturbance is inferior to that of acceleration. Most importantly, the benefit of using the proposed QZS isolators can be easily seen from a comparison between responses of the original transport incubator (Fig. 11) and the modified one (Fig. 12).

Fig. 12 Responses of the modified transport incubator (2-DOF lumped-mass model in Fig. 6b). (a) Displacements and (b) accelerations of infant and ambulance floor, when .

- Conclusions

A promising solution to vibration isolation in neonatal transport is proposed. Quasi-zero-stiffness (QZS) isolators are designed to attenuate transmission of disturbances from ambulance floor to infant in the incubator.  The lumped-mass dynamic model of transport incubator is established, and vibration-isolation performances under harmonic and random disturbances are evaluated in terms of displacement transmissibility and acceleration transmissibility.

The comparison of transmissibility between the original transport incubator and the modified one reveals the effectiveness of the proposed QZS vibration isolation system. Magnification of vibration is observed in the original transport incubator under a random disturbance, but more than 70% attenuation of Root Mean Square (RMS) acceleration and 50% attenuation of RMS displacement in the modified one. A comparatively heavy damping is needed to completely avoid the jump phenomenon and suppress resonance, and thus keep the QZS vibration-isolation system effective from an ultra-low frequency. The stretcher is suggested to be fixed rigidly on the ambulance floor during the neonatal transport, due to its negative effects on vibration isolation performance. The future work is to fabricate a prototype of the modified transport incubator and carry out experimental tests to verify its effectiveness in practice.

# Acknowledgments

This research work was supported by National Natural Science Foundation of China (11572116), Natural Science Foundation of Hunan Province (2016JJ3036), and Fundamental Research Funds for the Central Universities. This work is carried out during the first author’s visit to the University of Liverpool.

# Reference

ABAQUS V6.14 (2014) ABAQUS Analysis User's Manual. ABAQUS Inc.: Providence, RI, USA.

Bailey-Van Kuren M and Shukla A (2005) System design for isolation of a neonatal transport unit using passive and semi-active control strategies. Journal of Sound and Vibration 286: 382-394.

Bailey-Van Kuren M and Shukla A (2007) Isolation device for shock reduction in a neonatal transport apparatus, US Patents.

Beale R (2006) Patient transport method and apparatus. US Patents.

Browning J, Walding D, Klasen J and David Y (2008) Vibration Issues of Neonatal Incubators During In-Hospital Transport. Journal of Clinical Engineering 33(2): 74-77.

Carrella A, Brennan MJ, Waters TP and Shin K (2008) On the design of a high-static-low-dynamic stiffness isolator using linear mechanical springs and magnets. Journal of Sound and Vibration 315(3): 712-720.

Clark JG, Williams JD, Hood WB Jr and Murray RH (1967) Initial cardiovascular response to low frequency whole body vibration in humans and animals. Aerospace Medicine 38(5): 464-7.

Chen S, Wang Q and Wang T (2011) Hydroxy-terminated liquid nitrile rubber modified castor oil based polyurethane/epoxy IPN composites: Damping, thermal and mechanical properties. Polymer Testing 30(7): 726-731.

Draeger (2016) TI500 Globe-Trotter. Available at: http://www.draeger.com/sites/en_uk/Pages/Hospital/TI500-Globe-Trotter.aspx  (accessed 13 July, 2016)

Floyd WN, Broderson AB and Goodno JF (1973) Effect of whole-body vibration on peripheral nerve conduction time in the rhesus monkey.  Aerospace Medicine 44(3): 281-5.

Gajendragadkar G, Boyd JA, Potter DW, Mellen BG, Hahn GD and Shenai JP (2000) Mechanical vibration in neonatal transport: a randomized study of different mattresses. Journal of Perinatology 20(5): 307-10.

ISO-2631-1:1997 (1997) Mechanical vibration and shock - Evaluation of human exposure to whole-body vibration - part 1: general requirements.

Landstrom U and Lofstedt P (1987) Noise, vibration and changes in wakefulness during helicopter flight.  Aviation, Space, and Environmental Medicine 58(2): 109-18.

Le TD and Ahn KK (2011) A vibration isolation system in low frequency excitation region using negative stiffness structure for vehicle seat. Journal of Sound and Vibration 330: 6311-6335.

Osada Y and Kajiwar K (2001) Gels Handbook, Four-Volume Set, 1st Edition. San Diego: Academic Press.

Platus D (1991) Negative-stiffness-mechanism vibration isolation systems. In: Proceedings of the SPIE's International Symposium on Vibration control in Microelectronics, Optics and Metrology , San Jose, CA, USA, 1 November 1991, pp.44-54.

Prehn J, McEwen I, Jeffries L, Jones M, Daniels T, Goshorn E and Marx C (2015) Decreasing sound and vibration during ground transport of infants with very low birth weight. Journal of Perinatology 35(2): 110-4.

Sabota P and Aghili R (2014) Infant care transport device with shock and vibration system. US Patents.

Sallee W, Bentley A, Walding D and Christofi C (2016) Neonatal Transport Incubator: Vibration Identification, Ranking, and Attenuation—A Novel Approach to Patient Tray Stabilization. Journal of Clinical Engineering 41(2): 101-105.

Shah S, Rothberger A, Caprio M, Mally P and Hendricks-Munoz K (2008) Quantification of impulse experienced by neonates during inter- and intra-hospital transport measured by biophysical accelerometer. Journal of Perinatal Medicine 36(1): 87-92.

Shaw AD, Neild SA and Wagg DJ (2013) A nonlinear spring mechanism incorporating a bistable composite plate for vibration isolation. Journal of Sound and Vibration 332(24): 6265-6275.

Shenai JP, Johnson GE and Varney RV (1981) Mechanical Vibration in Neonatal Transport, Pediatrics 68(1): 55-57.

Sherwood HB, Donze A and Giebe J (1994) Mechanical Vibration in Ambulance Transport, Journal of Obstetric, Gynecologic & Neonatal Nursing 23(6): 457-463.

Sun X, Jing X, Xu J and Cheng L (2014) Vibration isolation via a scissor-like structured platform, Journal of Sound and Vibration 333(9): 2404-2420.

Wu W, Chen X and Shan Y (2014) Analysis and experiment of a vibration isolator using a novel magnetic spring with negative stiffness. Journal of Sound and Vibration 333(13): 2958-2970.

Xu D, Yu Q, Zhou J and Bishop S (2013) Theoretical and experimental analyses of a nonlinear magnetic vibration isolator with quasi-zero-stiffness characteristic. Journal of Sound and Vibration 332(14): 3377-3389.

Zheng Y, Zhang X, Luo Y, Yan B and Ma C (2016) Design and experiment of a high-static–low-dynamic stiffness isolator using a negative stiffness magnetic spring, Journal of Sound and Vibration 360: 31-52.

Zhong W (1998) Ferromagnetic Theory [in Chinese]. Beijing: China Science Press.

Zhou J, Wang X, Xu D and Bishop S (2015) Nonlinear dynamic characteristics of a quasi-zero stiffness vibration isolator with cam-roller-spring mechanisms, Journal of Sound and Vibration 346: 53-69.

Corresponding author.  College of Mechanical and Vehicle Engineering, Hunan University, Changsha 410082, PR China.

E-mail address: jxizhou@hnu.edu.cn (Jiaxi Zhou).