The Role of Gas in Determining Image Quality and Resolution During In Situ Scanning Transmission Electron Microscopy Experiments

Yuanyuan Zhu[a]*, and Nigel D. Browning[a, b]}TEM image quantification

fo

]}]]]]]]]][fnd;jfKf

}TEM image quantification

fo

]}]]]]

Abstract: As gas-solid heterogeneous catalytic reactions are molecular in nature, a full mechanistic understanding of the process requires atomic scale characterization under realistic operating conditions. While atomic resolution imaging has become a routine in modern high-vacuum (scanning) transmission electron microscopy ((S)TEM), both image quality and resolution nominally degrade when reaction gases are introduced. In this work, we systematically assess the effects of different gases at various pressures on the quality and resolution of images obtained at room temperature in the annular dark field STEM imaging mode using a differentially pumped (DP) gas cell. This imaging mode is largely free from inelastic scattering effects induced by the presence of gases and retains good imaging properties over a wide range of gas mass/pressures. We demonstrate the application of the ESTEM with atomic resolution images of a complex oxide alkane oxidation catalyst MoVNbTeOx (M1) immersed in light and heavy gas environments.

Introduction

Directly monitoring heterogeneous catalysts under catalytic conditions is key to obtaining insights into their behavior. Currently, environmental (scanning) transmission electron microscopy (E(S)TEM) is the only characterization tool that can provide atomic-site specific observations of local catalyst surface structure during solid-gas reactions.[1] In this approach, a volume of pressurized gas is introduced and maintained tightly around a TEM sample through either a differential pumping mechanism (DP) or by a membrane windowed (MW) gas cell.[2] The use of in situ TEM has progressed from pioneering early work featuring self-built modifications to existing microscopes/holders[3] to the use of commercially manufactured instruments[4]. However, the gaseous environment (and its container) continues limiting the (S)TEM sensitivity and resolvability in characterizing the solid catalyst due to an increase in scattering events from the background gases/windows. To achieve better imaging sensitivity and resolvability in catalytic relevant gaseous environments, it is therefore crucial to understand how the gas molecules affect the scattering of fast electrons contributing to the atomic scale image.[5]

There are two basic atomic imaging modes in the (S)TEM, the conventional high-resolution transmission electron microscope (HRTEM) phase contrast image and the incoherent annular dark-field scanning transmission electron microscopy (ADF-STEM) image. Although environmental HRTEM (HR-ETEM mode) has been widely used for in situ observations of heterogeneous catalysts and provides quantitative atomic information with the aid of image simulations,[6] environmental ADF-STEM (ESTEM mode) has already started demonstrating promising advantages for imaging supported metallic nanocatalysts[7] with single-atom sensitivity,[8] and for offering directly interpretable images of atomic scale dynamics in complex oxide catalysts.[9] Additionally, analytical techniques including energy-dispersive X-ray spectroscopy (EDX) and electron energy loss spectroscopy (EELS) in ADF-STEM mode (i.e. STEM-EDS and STEM-EELS) can also provide localized compositional and electronic information.[7b, 10] Despite these potential advantages of ESTEM the effect of the gas environment on the image quality has so far only been discussed for the HR-ETEM imaging mode.[5a, 11]

In this work, we conduct the first experimental evaluation on the quality and resolution of the DP-ESTEM imaging mode (in a non-aberration-corrected STEM) under the influence of various gas environments at different pressures. One of the main constrains of DP-ESTEM is the lower DP aperture which limits the ADF detector outer acceptance angle. We examine this issue using a uniform solid sample – the standard lacy amorphous carbon, whose low atomic weight (Z) renders it as a touchstone for evaluating the gas effects on ESTEM contrast and on image signal-to-noise ratio (SNR), with different ADF collection ranges. Since catalysts and particularly their atomic details in gas environments are usually the main interest of in situ microscopy observations, we also show how the gases affect the ESTEM SNR and resolution in atomic resolution images of crystalline MoVNbTeOx (M1) catalysts. All evaluations were conducted at room temperature, for high temperature imaging also depends on the stability of the heating stage used, which is beyond the scope of this study.

Results and Discussion

Effect of gas on ESTEM image quality for amorphous carbon (i.e. at low spatial resolution)

Figure 1a-d shows experimental DP-ESTEM images of regions with and without the solid carbon sample, in four commonly used gases at pressures from high vacuum (2x10-7 mbar) to 10 mbar.

Figure 1. Experimental ESTEM images of gas background and amorphous lacy carbon from high vacuum to 10 mbar in a) He, b) N2, c) H2 and d) Ar. All images are under the same magnification, and the scale bar is 200 nm. e) Intensity line profiles in raw counts from gas background to the carbon sample in N2 at different pressures. The measurement of Igas, Icarbon and ΔIcarbon are defined. f) Normalized experimental carbon intensity to corresponding Igas, showing the carbon contrast change against the N2 background at different pressures. The calculations of Contrastcarbon and SNRcarbon are defined. g) Normalized experimental ESTEM Contrast of carbon ((Icarbon-Igas)/Igas) and f) SNR of carbon in He, H2, N2 and Ar gas environments as a function of pressure. The Rose criterion (SNR > 5) is marked to show the preserved high SNR of carbon in heavy gas up to 10 mbar.

Considering that the DP gas cell apertures (esp. the ones below the sample) confine the scattering angle of exit electrons and impose a 71 mrad cut-off, we choose an ADF detector inner semi-angle of ~ 21 mrad to provide an adequate collection range of 21-71 mrad (i.e. a low-angle annular dark field (LAADF) with intensity ~ Z(, (: 1.0 ~ 1.3[12]). It is important to note that a low probe current of < 3 pA was used throughout this work to reflect the practical ESTEM conditions usually used for imaging beam-sensitive catalysts (for dose and dose-rate definitions see experimental section). As shown in, for example, the intensity line profiles of the pressure series in N2 gas (Figure 1e), the gas environment leads to an increasing intensity of the gas background (and of the solid sample as the gas is all around the sample); meanwhile, the noise level in the linescans also increases slightly along with N2 pressure (SI section 1).

More importantly, when the N2 pressure series is normalized to the corresponding gas background intensity (Igas), in Figure 1f, the contrast of the carbon decreases as the pressure increases. We define the ESTEM contrast of the lacy carbon as the difference between the intensity of carbon and gas background verses the gas intensity (Weber contrast): (Icarbon-Igas)/Igas. This ESTEM carbon contrast, when normalized to the initial contrast in the vacuum (Figure 1g) decreases by about ~ 55% in N2 and Ar at 10 mbar. This means that the “useful signal” that is from a sample is degrading as the gas pressure/mass increases. To examine to which extend the presence of gases influence the “image signal of a sample”, here, we define the signal/noise ratio (SNR) as the intensity from the carbon (Icarbon-Igas) divided by the standard deviation of the sample raw counts ΔIcarbon. Thus, the higher the SNR (the values in Figure 1h at low gas pressures), the better the imaging ability to observe the structural and chemical features of a sample. In contrast, as the gas pressure/mass increases the net carbon intensity starts to become comparable with its noise level, inducing uncertainties

Figure 2. Theoretically calculated cross sections of elastically and inelastically scattered electrons at different scattering angular ranges as a function of the atomic number of the object the electrons (with an incident energy of 300 kV) interact with. Different types of gases are marked based on their atomic number. For Titan DP environmental microscope, the scattered electrons collected in the ETEM imaging mode is 0-71 mrad (assume no objective aperture used), and an annular range of 21-71 mrad was collected by the LAADF-ESTEM imaging mode used in this work.

Figure 3. Experimental high resolution ESTEM images of the M1 catalyst in [001] zone axis from high vacuum to 10 mbar in a) He, and b) N2. Scale bar is 10 nm. c) The quantification of atomic ESTEM image SNR of M1 using the fast Fourier transform (FFT) measurement (details see experimental). d) The FFT measured SNR of the M1 experimental ESTEM images in He and N2 gas environments as a function of pressure.

in the observation of the solid sample. The SNR is found more 
strongly confined by the gas environment when the ESTEM images are collected under a medium-angle annular (MAADF) setting (SI section 2). However, for the LAADF mode with a relatively large collection range, the ESTEM SNR of the light carbon remains above the Rose criterion[13] in 10 mbar Ar. Thus, for a given catalytic sample (often a fixed Z-number) inner ADF collection angle can be used to leverage heavy gas and high pressure environment with controlled electron dose/dose-rate in DP-ESTEM, without introducing contrast reversals (even for light carbon) while maintaining a good sample contrast and SNR.

Considering that the STEM ADF detector is an electron-counting device, the raw image intensity faithfully counts the forward scattered electrons collected by the detector. The observed increases in ESTEM intensity and more importantly in the noise level from the area of the carbon suggest that there are additional “unwanted” electrons scattered by the gas molecules onto the ADF detector. These “extra electrons” cause the changes in the intensity as well as in the noise level in environmental images. To quantitatively assess to which extent this gas-induced scattering affects the ETEM and ESTEM imaging modes, we calculated the cross sections of elastic and inelastic scattering by the gases under the two imaging settings (SI section 3). As shown in Figure 2, the ETEM imaging mode records almost all the extra inelastic noise from the gases at low scattering angles (open symbols); while the ESTEM mode excludes most of it (inelastic characteristic scattering angles of typical gases are < 0.25 mrad in Table S1). In the ESTEM mode, the relatively low total extra scattering (from the gas molecules) explains the high imaging sensitivity (i.e. SNR) in the sample. This means that the advantage of ADF-STEM mode in revealing catalytic nanoparticles on a thick support background[14] also applies to imaging the materials in gases (the gas volume and the windows for WM gas cells are now the background of the signal).

Effect of gas on ESTEM atomic resolution image quality for the M1 catalyst

Next, we investigate the gas effect on the atomically resolved ESTEM images. Because of its large lattice parameter and good stability in gas environment,[9] the crystalline M1 catalyst was employed for this evaluation. Figure 3a-b present high-resolution ESTEM pressure series of the M1 crystal in [001] projection in He and N2 environment, respectively. To further quantify the gas effect, we performed a fast Fourier transform (FFT) based signal analysis which is better suited for measuring atomically resolved lattice images. For example, Figure 3c shows the FFT-based SNR measurement on an M1 atomic ESTEM image (Details of this method are given in the experimental section). The results show, in Figure 3d, the SNRM1 is, as expected, higher in the light He than that of in the heavier N2 at the same pressure. At 10 mbar, the ESTEM SNRM1 drops to 60.6% in He, and to 42.0% in N2 under the low beam current condition in this work. In the case of the He environment, since the ESTEM resolution hardly decreases up to 10 mbar, the degradation of the SNRM1 in He must be mainly caused by post specimen scattering of the image-forming electrons. That is to say, as the electrons carrying the sample information trying to reach the ADF detector, they could be further scattered by the gas molecules in the post-sample gas pathway,[5a] and are redirected to a final scattering angle that is outside of the detector collection range, leading to a loss of signal from the solid sample. For the gases heavier than He, both pre-sample and post-sample gas scatterings may contribute to the degradation in ESTEM SNRsample.

Figure 4a-b present the approach and results of the ESTEM resolution evaluation over the M1 crystal in He and in N2 up to 10 mbar. The large unit cell and thus densely packed diffraction spots of the M1 crystal in the [001] projection provide high sensitivity for resolution estimation. By measuring the ESTEM resolution on multiple M1 particles at different gas pressure (SI section 6), we find that the He gas hardly affected the ESTEM resolution (the resolution drop < 0.02 Å at 10 mbar) while the presence of N2 leads to noticeable resolution degradation (> 0.40 Å at 10 mbar), under the low beam current imaging condition. To understand this difference in the ESTEM resolution dependence on the gas type, we studied the cause of STEM resolution degradation. It is known that the resolution of STEM is governed by the size of the converged probe.[15] In conventional microscopes, the integrity of the STEM probe is preserved due to the large electron mean free path under high vacuum; in the DP gas cell, prior to interacting with the sample the STEM probe has to pass through the pre-sample gas species that lead to extra scattering and possible diffuse of the originally sharp probe. If we assume that the gas species along the pre-sample pathway are acting collectively as a solid foil, the gas-induced STEM probe spreading can then be considered in a similar way to the beam broadening b caused by a (thick) sample. This probe broadening b can be estimated quantitatively based on Bothe’s multiple–scattering theory,[16]

(1)

Figure 4. a) Experimental ESTEM resolution estimation based on the FFT of an atomic M1 image. b) Theoretical and experimental ESTEM resolution of M1 in He and N2 up to 10 mbar. c) Schematic illustration of pre-sample gas introduced STEM probe spreading.

where the ρ, W and Z are the mass density, the mean molar mass and the mean atomic number of a chosen solid, E is the microscope acceleration voltage in eV and E0 is 1 eV. It shows that b increases with the 3/2 power of the sample thickness.

Considering that silicon nitride (Si3N4) is commonly used as the window material for the membrane windowed (MW) environmental gas cell, we chose it as the gas-equivalent solid foil. Based on the idea gas law, at a given pressure (P) and temperature (T) the reversed molar volume of a gas with a thickness of Lgas is equal to the reversed molar volume of Si3N4 with a thickness of LSi3N4,

(2)

where R is the gas constant, W is 20.0 g/mol and ρ is 3.2×106 g/m3 for Si3N4. However, the above equation does not specify the type of the gas, i.e. any 10 mbar gases at 300 K with a gas pathway of 2.7 mm are equivalent to 6.75 nm Si3N4. Obviously, this simple estimation is too generalized to any 10-mbar gases regardless of their type and thus can’t explain the difference in ESTEM resolution observed in He and in N2. To take the nature of the gas into consideration, we revise the above equation by weighting both sides with the total scattering cross section σ and distinguishing gases by including the number of atoms in a gas molecule M,

(3)

With this modification, the equivalent Si3N4 thicknesses of He and N2 are 0.51 nm and 21.29 nm, respectively at 10 mbar (Table S2). Thus, the theoretical probe broadening b as well as the corresponding ESTEM resolution reduction ΔR were estimated. The theoretically estimated ESTEM resolution drops are plotted as solid lines in Figure 4b. It shows that at 10 mbar, ESTEM resolution change is negligible in the He gas; while 2.7 mm N2 leads to about ~ 0.25 Å degradation in ESTEM resolution in theory. Note that the apparent probe broadening b should be added to the original probe size in quadrature to reflect the Gaussian-distribution nature of the STEM probe intensity.

Thus, as schematically illustrated in Figure 4c, after taking the gas type into account the above theoretical estimation explains qualitatively the observed difference in the trend of the ESTEM resolution up to a higher pressure between the light and heavy gas environment. It is worth mentioning that the pessimistic ESTEM resolution measured in N2 (blue triangles in Figure 4b) was under the low beam current imaging condition used throughout this work. This low-dose resolution can be boosted by adjusting the STEM imaging conditions if the sample does not damage and there are no kinetic changes in reaction caused by the elevated current (details will be discussed in the next section).

Comments on Electron Dose and Dose Rate

To examine if the low beam current imaging conditions used in this work become limiting factors for ESTEM image quality and resolution with the presence of gas, we tested the above ESTEM SNRs, contrast and resolution at different electron dose and dose rates, respectively, in 10 mbar N2 environment. As shown in Figure 5, both of the SNRs of the amorphous carbon and of the M1 sample, and the ESTEM resolution increase along with the increase of total accumulated electron dose per image frame; whereas, such correlations are not observed when varying the dose rate (Figure S7). In particular, for the atomically resolved ESTEM images of the M1 catalyst in 10 mbar N2, by increasing the electron dose from the pre-set low dose (8.24 x 102 e/Å2) to a relatively higher dose of 6.57 x 103 e/Å2, the SNRM1 is increases from 7.21 to 15.20 (the orange square in Figure 3d), and the resolution drop is mitigated to ~ 0.20 Å (Figure 4b), agreeing well with the theoretical resolution estimation. This suggests that both ESTEM SNR and resolution depend on the accumulated electron dose in an image, which is fundamentally different from the dose-rate dependence found for the ETEM resolution.[11a, 11b] On the other hand, the ESTEM contrast of the amorphous carbon remains largely constant

Figure 5. Experimentally measured electron dose dependent of a) the ESTEM SNR and contrast of carbon, and b) the SNR and ESTEM resolution (Figure S6) over the M1 catalyst, in 10 mbar N2.

regardless of the electron dose and dose rate used for imaging (note that the fluctuations in contrast at extremely low dose/dose-rate are caused by the high uncertainties in determining the contrast). This suggests that the “Z-contrast” nature of the STEM imaging mode, and thus the elemental indications of the STEM intensity, is still valid for ESTEM imaging with the consideration for environmental contribution.

Unlike the convoluted relationship between the TEM beam and gas species,[5] a relatively simple picture might be possible to elucidate the gas effects in the STEM imaging mode. One such attempt is schematically illustrated in Figure 6a. In the post-sample region at the bottom half of the gas-cell chamber, the STEM image-forming electrons carrying the information of the sample are generally inelastic in nature. The interactions between these electrons (i.e. sample signal) and the gas species (and/or bottom silicon nitride window) mainly lead to a diffusion of the STEM signal by scattering (some of) the image-forming electrons off the finite ADF detector. Thus, it results in losses in SNR and in resolution due to the lack of signal. By increasing the electron dose via a longer dwell time for example, the signal level can be retained and then recovers the ESTEM SNR and resolution. In another word, if a catalyst sample can tolerate accumulated electron dose, the effects of gas species in the post-sample region can be largely alleviated. For example,

Figure 6. a) Schematic illustration of the pre- and post-sample gas effects on ESTEM probe and signal. b) Theoretical estimated ESTEM resolution as a function of silicon nitride thickness (Table S3).

excellent atomic ESTEM SNR was observed using electron dose ~ 104 e/Å2.[7b]

With the ESTEM signal diffusion retrieved via accumulating more electron doses, we can then propose the theoretically obtainable ESTEM image resolution for both DP-ESTEM and WM-ESTEM. Based on the multiple-scattering theory, in Figure 6b, the estimated resolution for an original STEM probe size of 0.78 Å (e.g. JEM-ARM 200F, probe-corrected), of 1.36 Å (e.g. Titan G2 300kV) and of 1.90 Å (e.g. Tecnai G2 F30 S-Twin) are plotted as a function of the thickness of the gas-equivalent silicon nitride (window). In general, the ESTEM probe broadening accelerates as the gas mass/pressure (or the pre-sample window thickness) becomes greater. For the DP gas cell (used in this work), the effect of 10 mbar N2 on the ESTEM resolution is similar to the best obtainable resolution of a WM gas cell with a 20-nm-thick pre-sample Si3N4 window (when assuming the sample is attached to the top inside of the upper window). In a similar sense, the upper limit of 20 mbar N2 in a DP gas cell (of 2.7 mm pre-sample gas pathway) is equivalent to a WM cell of ~ 43 nm Si3N4 window (not shown here). On the other hand, as indicated by the comparison between a 10 nm and a 50 nm Si3N4 pre-sample window, the achievable ESTEM resolution for WM gas cells depends strongly on the window thickness (and for that matter, also the sample position). To eliminate the probe broadening effects, a Si3N4 window thinner than < 6 nm (or the N2 in 2.7-mm gas-pathway is lower than ~ 3 mbar) is required. For a practical ESTEM probe of < 2.0 Å applicable for most of metallic and metal oxide catalysts, advanced objective lens is encouraged for a DP-ESTEM and probe correction is required for a WM-ESTEM.

Conclusions

Catalysis has benefited greatly from atomic-level imaging in modern (S)TEM, and in turn the kinetic nature of catalysis has also shaped the development of new capabilities in (S)TEM. The recent advancements in situ environmental (S)TEM have further opened up unprecedented opportunities for obtaining experimental insights into the working structure of heterogeneous catalysts in technologically relevant environments. Meanwhile, the presence of catalytic environment such as pressurized gases poses new challenges to the (S)TEM sensitivity and resolvability and to our understanding of gas-electron interactions. In this work, we investigated both experimentally and theoretically the gas environment effects on image quality (in terms of sample contrast and SNR) and resolution in the directly interpretable ADF-STEM imaging mode in DP-ESTEM. The conclusions and guidelines identified are as follows:

- The preconceived constrain imposed by the lower gas aperture in DP-ESTEM can be alleviated by loosening the inner ADF collection angle, without introducing image contrast reversals.

- The absence of the extra inelastic noise due to the annular collection mode of the ADF-STEM imaging allows the ESTEM imaging mode to be free of intensity loss, and to maintain a good sensitivity in imaging the solid sample (sample SNR) with the presence of gases. Unlike the ETEM mode, ESTEM does not require a monochromator to achieve high quality imaging.

- The ESTEM sample contrast, owing to the incoherent nature of the ADF-STEM imaging mode, remains being correlated to the (mean) atomic number of the solid sample with the consideration of the elastic contribution from the gas environment.

- The obtainable ESTEM resolution is determined mainly by the probe broadening effects induced by the pre-sample gas and electron beam interactions. Our theoretical estimation on the resolution shows that probe-corrector is a necessity for high-resolution ESTEM imaging.

- The practical ESTEM sample SNR and resolution are electron dose dependent, due to the signal diffusion effect caused by the gas species in the post-sample region. In practice, the property that the SNR and resolution both increase along with the accumulation electron dose is beneficial for achieving high sensitivity at the same time high-resolution ESTEM observations.

The ultimate challenge in imaging working catalysts in gas-solid heterogeneous catalysis is to be able to directly visualize the gas adsorbent at a catalyst surface with the reaction rate relevant high temporal resolution. Phase contrast in conventional HRTEM has been used to image light materials (i.e. weak-phase objects) with the additional help of either phase plate[17] or using off-axis holography[18]. And although the frame time for one TEM image is relatively short than a typical STEM image, a HRTEM focal-series is required for atomic interpretation which takes a much longer time and then low in temporal resolution. Recent development in advanced STEM imaging has began to leverage light-element imaging by for example ptychography,[19] and to accelerate scanning speed without compromising image SNR by novel scan sampling[20]. The combinations of these recent developments with the current high quality atomic ESTEM discussed in this work, makes it a promising in situ tool-box for investigating catalytic dynamics.

Experimental Section

Sample preparation

Standard TEM lacey carbon Cu grids (Ted-Pella) were used as it is for ESTEM imaging of amorphous carbon. High purity crystalline MoVTeNb mixed oxide catalyst (M1-phase) was obtained by hydrothermal synthesis.[9] TEM samples were prepared by crushing the M1 catalyst into fine powder, and dry loading on ultrathin carbon film on lacey carbon support film Au grids (Ted-Pella). M1 samples were plasma cleaned 10 to 15 s prior to ESTEM imaging.

Environmental scanning transmission electron microscopy imaging

We employed a dedicated ETEM (FEI Titan 80/300) with differentially pumped gas cell system (objective lens pole pieces separation ~ 5.4 mm) housed in EMSL user facility at PNNL. All ESTEM images shown in this work were acquired at 300kV with a convergence semi-angle of 9.9 mrad without probe aberration correction. A low-angle annular dark field (LAADF) STEM mode with a collection range of 21 – 71 mrad (upper collection angle is limited by the DP gas cell aperture) was used throughout this work. A higher ADF detector inner semi-angle of 48 mrad (MAADF-STEM) was also tested with the presence of gases and details on the justification of using the LAADF-ESTEM imaging mode see SI section 2. This microscope equipped with a field-emission gun produces stable and bright electron probe. To determine the accurate electron current on sample, we conducted a faraday cup measurement using an analytical holder (Gatan, Inc.) to calibrate the microscope screen dose-mete especially for the low current density region of < 3 pA. To reproduce the low electron dose-rate and low accumulated dose imaging conditions that often used for imaging delicate catalysts, a low beam current was kept around1.5-3 pA and a short dwell time of 4 μs/pixel was used for collecting all the ESTEM images (unless stated otherwise). The image magnification (and hence dose-rate) was chosen based on the size of the sample feature imaged. In particularly, for the amorphous carbon, we used a magnification of x115 k corresponding to 3.34 Å/pixel, dose rate 0.26 e/Å2s and dose per frame 4.50 e/Å2; the common magnification used for atomically resolved M1 catalyst is x1800 k corresponding to 0.22 Å/pixel, dose rate 4.79 x 101 e/Å2s and dose per frame 8.24 x 102 e/Å2 (unless stated otherwise). The total dose was kept under the damage threshold for each M1 particle imaged.[9]

Contamination from hydrocarbon build up in the ETEM column can impose difficulties during STEM imaging especially for a multi-user instrument. In this work, we found that the built-in plasma cleaner in the ETEM helped alleviate this issue. Prior to each ESTEM session, the plasma clean procedure was performed at 14 W (energy) for 10 hours followed by 4 hours purging with nitrogen. External plasma cleaning of the TEM sample is also encouraged to further eliminate contamination. During the ESTEM experiment, the back-pressure around the sample was pumped down to < 2 x 10-7 mbar after inserting the sample holder to set up a clean background. Ultra-pure He, H2, N2 and Ar was introduced around the sample to 0.1, 1, 5 and 10 (± 5%) mbar, and flowed for about 10 min every time after reaching the targeting pressure before imaging. The optimal STEM imaging defocus was achieved by maximizing the contrast.

ESTEM image quantification

It is essential to read out the raw electron counts from an ESTEM image instead of the apparent image intensity, which varies depending on the brightness/contrast (B/C) setting. To avoid this, we used the original 16-bit .ser files or converting them to 16-bit (not 8-bit) .tiff image files (so they can be read in DigitalMicrograph) for performing image intensity line profile. However, the ESTEM images shown in the figures in this paper were converted to 8-bit displays with a similar B/C setting to provide a reasonable demonstration of the gases effect.

For the ESTEM images of the gas background and amorphous carbon, since both are relatively uniform applying the line profiling and extracting raw count intensity is straightforward (Figure 1). However, for the atomically resolved ESTEM images, line profile measurement could introduce ambiguity. For example, SNR measured by line profiling are different depends on if it is along the slow or the fast scan direction.[21] Here, we adopted a Matlab script using the fast Fourier transform (FFT) magnitudes measurement developed by Colin Ophus at Lawrence Berkeley National Laboratory. This script locates the Bragg peaks on the FFT pattern of an ESTEM lattice image. The sum of the Bragg peak intensities is the “useful signal” (amplitude squared), the sum of noise pixel intensities is the “noise”, and the ratio of the two gives SNR. This method provides a relative SNR but the measurement is robust and highly reproducible. Note that in order to compare ESTEM images with different pixel sampling (thus different size of k frequency space in FFT), we defined a unified k frequency cutoff by applying a soft Gaussian low pass filters. By summing the signals and noise within the same size of k-space for different magnifications, we can then compare SNR across different reciprocal pixel size.

Acknowledgements

This research is part of the Chemical Imaging Initiative and a seed project conducted under the Laboratory Directed Research and Development Program at Pacific Northwest National Laboratory (PNNL). PNNL, a multiprogram national laboratory, is operated by Battelle for the Department of Energy under Contract DE-AC05-76RLO1830. A portion of the research was performed using Environmental Molecular Sciences Laboratory (EMSL), a national scientific user facility sponsored by the Department of Energy's Office of Biological and Environmental Research and located at Pacific Northwest National Laboratory. Special thanks to Dr. Colin Ophus at Molecular Foundary, LBNL, for his generous help and providing the FFT magnitude Matlab script. The authors thank Dr. Libor Kovarik (EMSL, PNNL) for technical support on the ESTEM, and we also thank Daniel Melzer and Dr. Maricruz Sanchez-Sanchez at Technical University of Munich for providing the M1 catalysts.

Keywords: Gas-solid reaction • Environmental scanning transmission electron microscopy • Gas atmosphere • Sample SNR • Image resolution

[1]	a) J. M. Thomas, P. A. Midgley, ChemCatChem 2010, 2(7), 783-798; b) R. Sharma, Micron 2012, 43(11), 1147-1155; c) P. A. Crozier, T. W. Hansen, MRS Bull. 2015, 40(1), 38-45; d) E. D. Boyes, P. L. Gai, MRS Bull. 2015, 40(7), 600-605; e) S. Helveg, C. F. Kisielowski, J. R. Jinschek, P. Specht, G. Yuan, H. Frei, Micron 2015, 68, 176-185.

[2]	L. Marton, Bull. Acad. r. Belg. Cl. Sci. 1935, 21.

[3]	a) H. G. Heide, J. Cell Biol. 1962, 13(1), 147; b) E. J. Gallegos, Rev. Sci. Instrum. 1964, 35(9), 1123; c) H. Hashimoto, T. Naiki, Japanese Journal of Applied Physics 1968, 7(8), 946-+; d) P. R. Swann, High voltage microscope studies of environmental reactions, Univ. Calif., 1972, pp. 878-904; e) H. M. Flower, Journal of Microscopy-Oxford 1973, 97(JAN-M), 171-190; f) N. M. Rodriguez, S. G. Oh, W. B. Downs, P. Pattabiraman, R. T. K. Baker, Rev. Sci. Instrum. 1990, 61(7), 1863-1868; g) R. Sharma, K. Weiss, M. McKelvy, W. Glaunsinger, in Fifty-Second Annual Meeting - Microscopy Society of America/Twenty-Ninth Annual Meeting - Microbeam Analysis Society, Proceedings (Eds.: G. W. Bailey, A. J. GarrattReed), 1994, pp. 494-495; h) E. D. Boyes, P. L. Gai, Ultramicroscopy 1997, 67(1-4), 219-232.

[4]	a) H. Yoshida, S. Takeda, Phys. Rev. B. 2005, 72; b) J. F. Creemer, S. Helveg, G. H. Hoveling, S. Ullmann, A. M. Molenbroek, P. M. Sarro, H. W. Zandbergen, Ultramicroscopy 2008, 108(9), 993-998; c) P. L. Gai, E. D. Boyes, Microsc. Res. Tech. 2009, 72(3), 153-164; d) T. W. Hansen, J. B. Wagner, R. E. Dunin-Borkowski, Mater. Sci. Technol. 2010, 26.

[5]	a) M. Suzuki, T. Yaguchi, X. F. Zhang, Microscopy (Oxford, U. K.) 2013, 62(4), 437-450; b) J. B. Wagner, M. Belaggia, in Controlled atmosphere transmission electron microscopy: principles and practice (Eds.: T. W. Hansen, J. B. Wagner), Springer International Publishing, 2016.

[6]	a) P. L. Hansen, J. B. Wagner, S. Helveg, J. R. Rostrup-Nielsen, B. S. Clausen, H. Topsoe, Science 2002, 295(5562), 2053-2055; b) H. Yoshida, Y. Kuwauchi, J. R. Jinschek, K. J. Sun, S. Tanaka, M. Kohyama, S. Shimada, M. Haruta, S. Takeda, Science 2012, 335(6066), 317-319.

[7]	a) P. Li, J. Liu, N. Nag, P. A. Crozier, Applied Catalysis a-General 2006, 307(2), 212-221; b) S. Zhang, P. N. Plessow, J. J. Willis, S. Dai, M. J. Xu, G. W. Graham, M. Cargnello, F. Abild-Pedersen, X. Q. Pan, Nano Lett. 2016, 16(7), 4528-4534.

[8]	P. L. Gai, L. Lari, M. R. Ward, E. D. Boyes, Chem. Phys. Lett. 2014, 592, 355-359.

[9]	Y. Zhu, P. V. Sushko, E. Jensen, L. Kovarik, D. Melzer, C. Ophus, M. Sanchez-Sanchez, J. Lercher, N. Browning, 2017, In review.

[10]	P. A. Crozier, in Scanning Transmission Electron Microscopy (Eds.: S. J. Pennycook, P. D. Nellist), Springer New York, 2011, pp. pp 537-582.

[11]	a) J. R. Jinschek, S. Helveg, Micron 2012, 43(11), 1156-1168; b) A. N. Bright, K. Yoshida, N. Tanaka, Ultramicroscopy 2013, 124, 46-51; c) H. L. Xin, K. Niu, D. H. Alsem, H. Zheng, Microsc. Microanal. 2013, 19(6), 1558-1568; d) M. Ek, S. P. F. Jespersen, C. D. Damsgaard, S. Helveg, Advanced Structural and Chemical Imaging 2016, 2(1), 1-8.

[12]	a) M. M. J. Treacy, Journal De Microscopie Et De Spectroscopie Electroniques 1982, 7(5), 511-523; b) Z. W. Wang, Z. Y. Li, S. J. Park, A. Abdela, D. Tang, R. E. Palmer, Physical Review B 2011, 84(7).

[13]	A. Rose, in Advances in Electronics and Electron Physics, Vol. Volume 1 (Ed.: L. Marton), Academic Press, 1948, pp. 131-166.

[14]	V. Ortalan, A. Uzun, B. C. Gates, N. D. Browning, Nat. Nanotechnol. 2010, 5(12), 843-847.

[15]	N. D. Browning, M. F. Chisholm, S. J. Pennycook, Nature 1993, 366(6451), 143-146.

[16]	L. Reimer, Transmission electron microscopy, Springer:Heidelberg, 1984.

[17]	R. Danev, H. Okawara, N. Usuda, K. Kametani, K. Nagayama, J. Biol. Phys. 2002, 28(4), 627-635.

[18]	M. Linck, B. Freitag, S. Kujawa, M. Lehmann, T. Niermann, Ultramicroscopy 2012, 116, 13-23.

[19]	H. Yang, R. N. Rutte, L. Jones, M. Simson, R. Sagawa, H. Ryll, M. Huth, T. J. Pennycook, M. L. H. Green, H. Soltau, Y. Kondo, B. G. Davis, P. D. Nellist, Nature Communications 2016, 7.

[20]	a) L. Kovarik, A. Stevens, A. Liyu, N. D. Browning, Appl. Phys. Lett. 2016, 109(16); b) T. Dahmen, M. Engstler, C. Pauly, P. Trampert, N. de Jonge, F. Mucklich, P. Slusallek, Scientific Reports 2016, 6.

[21]	Y. Zhu, C. Ophus, J. Ciston, H. Wang, Acta Mater. 2013, 61(15), 5646-5663.

Entry for the Table of Contents

Layout 1:

| FULL PAPER |
| --- |
| In situ environmental (scanning) transmission electron microscopy has opened up new opportunities for studying the atomic-level dynamics in heterogeneous catalysis. This work provides the first evaluation on the sensitivity and resolution of STEM imaging mode with the presence of gases. Exempted from gas-induced inelastic scattering, ESTEM offers superior sensitivity and predictable resolution in imaging solid catalysts under realistic reaction conditions. |  |  |  | Yuanyuan Zhu1★, Nigel D. Browning1, Page No. – Page No. The Role of Gas in Determining Image Quality and Resolution During In Situ Scanning Transmission Electron Microscopy Experiments |

# Text boxes

[a]	Dr. Y. Zhu, and Prof. Dr. N.D. Browning
Physical & Computational Science Directorate
Pacific Northwest National Laboratory
Richland, WA 99352 (USA)
E-mail: yuanyuan.zhu@pnnl.gov; nigel.browning@pnnl.gov

[b]       Prof. Dr. N.D. Browning

Department of Materials Science and Engineering

University of Washington,

Seattle, WA 98195 (USA)

Supporting information for this article is given via a link at the end of the document