**Determination of Corneal Biomechanical Behaviour In-vivo for Healthy Eyes Using CorVis ST Tonometry: Stress-Strain Index**

Ashkan Eliasy\*1; Kai-Jung Chen1; Riccardo Vinciguerra1, 2; Bernardo T. Lopes1, 5, 6; Ahmed Abass1; Paolo Vinciguerra3, 4; Renato Ambrósio Jr5, 6; Cynthia J. Roberts7; Ahmed Elsheikh1, 8, 9;

**Affiliation:**

1School of Engineering, University of Liverpool, Liverpool L69 3GH, UK

2St Paul’s Eye Unit, Royal Liverpool and Broadgreen University Hospital, Liverpool, UK

3Department of Biomedical Science-Humanitas University, Via Manzoni 56, Rozzano (MI) – Italy.

4Eye Center, Humanitas Clinical and Research Center, Via Manzoni 56, Rozzano (MI) – Italy.

5Rio de Janeiro Corneal Tomography and Biomechanics Study Group – Rio de Janeiro, Brazil

6Department of Ophthalmology, Federal University of the State of Rio de Janeiro – Rio de Janeiro, Brazil

7Department of Ophthalmology & Visual Science, Department of Biomedical Engineering, The Ohio State University – Columbus, OH, USA

8NIHR Biomedical Research Centre for Ophthalmology, Moorfields Eye Hospital NHS Foundation Trust and UCL Institute of Ophthalmology, UK

9 School of Biological Science and Biomedical Engineering, Beihang University, Beijing, China

**Keywords:** cornea; biomechanics; material properties; numerical modelling; finite element modelling

**Author for correspondence:**

Ashkan Eliasy

School of Engineering, University of Liverpool, Liverpool L69 3GH, UK

eliasy.ashkan@gmail.com

**Running Title:** IntroducingStress-Strain Index (SSI)

**Number of words:** 3500

**Abstract**

**Purpose:** This study aims to introduce and clinically validate a new algorithm that can determine the biomechanical properties of the human cornea in vivo.

**Methods:** A parametric study was conducted involving representative finite element models of human ocular globes with wide ranges of geometries and material biomechanical behaviour. The models were subjected to different levels of intraocular pressure (IOP) and the action of external air puff produced by a non-contact tonometer. Predictions of dynamic corneal response under air pressure were analysed to develop an algorithm that can predict the cornea’s material behaviour. The algorithm was assessed using clinical data obtained from 480 healthy participants where its predictions of material behaviour were tested against variations in central corneal thickness (CCT), IOP and age, and compared against those obtained in earlier studies on ex-vivo human ocular tissue.

**Results:** The algorithm produced a material stiffness parameter (Stress-Strain Index or SSI) that showed no significant correlation with both CCT (p> 0.05) and IOP (p> 0.05), but was significantly correlated with age (p< 0.01). The stiffness estimates and their variation with age were also significantly correlated (p< 0.01) with stiffness estimates obtained earlier in studies on ex-vivo human tissue.

**Conclusions:** The study introduced and validated a new method for estimating the in vivo biomechanical behavior of healthy corneal tissue. The method can aid optimization of procedures that interfere mechanically with the cornea such as refractive surgeries and introduction of corneal implants.

**Introduction**

The ability to determine corneal biomechanical properties in-vivo is of great clinical importance as it can help optimise several treatments and management procedures that interact or interfere mechanically with the eye. Examples include measurement of intraocular pressure (IOP) for effective glaucoma management (Kaushik et al., 2012;Elsheikh et al., 2015), refractive surgery planning (Roberts, 2002;Pepose et al., 2007), keratoconus risk profiling (Ortiz et al., 2007;Ambrosio et al., 2017a), optimisation or judging different protocols of collagen cross-linking treatments (Goldich et al., 2012), pre-op evaluation of refractive surgery re-treatment, selection of intracorneal ring implants and even design of soft contact lenses where the mechanical interaction between the lens and the anterior eye is currently not considered.

A main challenge in estimating the corneal biomechanical behaviour in vivo stems from the difficulty in separating the effects of this behaviour from those of the IOP on ocular response to mechanical stimuli. This challenge has made it difficult to produce accurate IOP estimates, that are free of the effects of corneal biomechanics (Liu and Roberts, 2005), and the same challenge exists in determining the tissue’s biomechanics that are free of the effects of IOP. Nevertheless, the compound nature of this challenge has meant that finding a solution for either IOP or corneal biomechanics would lead to a solution for the other problem.

What complicates matters further is that the stress-strain behaviour of biological tissue, including cornea and sclera, is nonlinear (Ethier et al., 2004;Elsheikh et al., 2007), and therefore the tangent modulus (Et) – a measure of material stiffness – does not have a constant value, but increases with stress and strain. This effectively means that as the IOP in the eye increases, the stress and strain to which the eye is subjected increases, causing a rise in the tangent modulus. Therefore, the problem is not only that the effects of IOP and corneal biomechanics on eye behaviour are difficult to separate; IOP also effects the immediate corneal stiffness.

A positive development towards achieving a solution to this problem was the introduction of the biomechanically-corrected IOP (bIOP) estimates based on the CorVis ST (OCULUS Optikgeräte GmbH; Wetzlar, Germany) output (Elsheikh et al., 2015). The bIOP algorithm was developed using a combination of numerical modelling, experimental and clinical validation (Elsheikh et al., 2010a;Joda et al., 2016), as well as corneal deformation parameters (measured by the CorVis ST) to reduce the effect of stiffness on IOP estimates (Eliasy et al., 2018). With the bIOP shown in earlier studies to be less correlated with the cornea’s stiffness parameters than both GAT and the uncorrected CorVis ST IOP (CVS-IOP) measurements (Chen et al., 2018), this study takes the next logic step in providing estimates of the material mechanical behaviour.

This step is taken in this study where the emphasis is on an algorithm that can provide an estimate of the whole stress-strain behaviour that would, in turn, enable determination of Et under any IOP, and would ultimately be suitable for use in numerical simulation exercises to exploit the benefits of material characterisation in clinical applications such as refractive surgery planning or cross-linking therapy optimisation.

*Hypothesis: The study was based on the hypothesis that a biomechanical CorVis index can be numerically developed and shown to be almost independent of CCT and IOP but maintained positive correlation with age in healthy patients.*

**Methods**

The study relied on numerical models of the full eye globe subjected to both IOP and the air pressure of the CorVis ST. The models enabled simulation of wide ranges of ocular topography, thickness profiles, IOP values and material behaviour trends that extend beyond those seen in ophthalmic practice or reported in the literature. The analysis resulted in predictions of corneal deformation and CorVis output parameters for each combination of the input parameters, and these predictions were used to develop an algorithm providing estimates of the tissue’s material behaviour as a function of the cornea’s geometric parameters, the IOP measurement and the CorVis output parameters. The algorithm was then validated by assessing the correlation between its material stiffness predictions and patient age in two clinical datasets, and against earlier results of inflation experiments on ex-vivo human eyes (Eliasy et al., 2018).

***Numerical Modelling***

Finite element models of full eye globes were developed by a bespoke ocular mesh-generator software tool (developed in house (Whitford et al., 2015)) and analysed using Abaqus 6.14 FE solver (Dassault Systèmes Simulia Corp., Providence, RI, USA), Figure 1. The models included 65712 six-noded, continuum C3D6H elements, connected by 65716 nodes, and organized in 25 cornea element rings and 124 sclera element rings, Figure 2.

Rigid-body motion of the models was prevented by restricting the equator nodes in the anterior-posterior direction, and the corneal apex in both the superior-inferior and temporal-nasal directions, Figure 2. The models had a fluid cavity filled with an incompressible fluid with a density of 1000 kg/m3 to simulate the aqueous and vitreous and their incompressible behaviour (Villamarin et al., 2012). IOP was applied and varied in the model through controlling the pressure in this internal fluid. This technique enabled the internal eye pressure to vary from the initial IOP according to the deformation experienced under the CorVis air pressure. At the start of the analysis, the stress-free form of each model, which corresponded with a state under IOP = 0 mmHg, was reached using an iterative process (Elsheikh et al., 2013) before applying IOP followed by the CorVis air pressure.

The eye model was divided into four regions incorporating the cornea, limbus, anterior sclera and posterior sclera, with different stress-strain behaviour patterns. Third-order, hyper-elastic Ogden models were used to represent the ocular tissue’s mechanical behaviour as obtained in previous experimental studies where correlation between stress-strain behaviour and age was reported (Elsheikh et al., 2010b;Geraghty et al., 2012). Moreover, scleral regional variation in stiffness, with a gradual reduction in stiffness from the limbus towards the posterior pole, was incorporated in the numerical models (Elsheikh et al., 2010a).

***CorVis Simulation***

The air puff of the CorVis was applied on ocular numerical models as per the results of experiments provided by the manufacturer and depicted in Figure 3 (Elsheikh et al., 2009). The results indicated a maximum air pressure of 95 mmHg at corneal apex, reducing away from the apex to a zero value at 4mm radius. Figure 3B shows the profile of pressure applied by the CorVis on the cornea, which starts with a 5 ms stage with relatively low pressure followed by a fast rise then fall of pressure within approximately a 22 ms period.

***Parametric study***

The numerical models were used in a parametric study that covered wide variations in IOP, geometry and material parameters. IOP varied between 10 and 30 mmHg (in steps of 5 mmHg), covering the values commonly seen in ophthalmic practice, while central corneal thickness (CCT) varied between 445 and 645 microns (in steps of 50 microns). These values of CCT covered and slightly extended beyond the ranges reported in clinical studies, while corneal curvature was fixed at 7.8mm (Dubbelman et al., 2002;Belin and Khachikian, 2006;Gilani et al., 2013). The peripheral corneal thickness (PCT) was assumed larger than CCT by 150 microns (Avitabile et al., 1997;Ambrosio et al., 2006) with a linear growth in thickness between the two, and in the sclera, the thickness varied linearly from PCT at the limbus, to 80% of PCT at the equator and 120% of PCT at the posterior pole, based on findings of an earlier experimental study (Elsheikh et al., 2010a). The optic nerve head was not simulated in the models as its effect on corneal behaviour was expected to be insignificant.

In order to consider variations in the tissue’s material properties, experimental stress-strain behaviour obtained in earlier studies (Elsheikh et al., 2010b;Geraghty et al., 2015) by the Biomechanical Engineering Group was assessed and found to follow the similar trends depicted in Figure 4b, rather than the intersecting trends shown in Figure 4a. This feature meant that different stress-strain curves could be obtained from each other while applying a simple stretching factor as a multiplier to all strain values. This factor, called in this study the Stress Strain Index, or SSI, was taken as 1.0 for the average experimental behaviour obtained for corneal tissue with age = 50 years (Elsheikh et al., 2010b). Higher values of SSI would then be indicative of higher tissue stiffness, and vice versa.

The average behaviour of 50 year old tissue was determined experimentally as (Elsheikh et al., 2010b):

σ = 1.26 × 10-3 × (e102.9 × ε -1) (Eq.1)

As this σ - ε behaviour is approximately exponential, the resulting Et - σ behaviour would be almost linear. This feature enabled making the changes in Stress-Strain Index (SSI) proportional to the changes in Et at any stress level as depicted in Figure 4c. The parametric study considered variations in SSI from 0.30 to 3.00, representing a range of stiffness from very soft to very stiff, respectively.

At the end of each simulation, the eye model’s deformation under IOP and CorVis ST pressure was recorded and used to predict values of the main CorVis corneal deformation parameters, including the highest concavity radius, maximum deflection, first applanation pressure and first applanation deflection (Nemeth et al., 2013;Roberts et al., 2017).

***Algorithm to estimate SSI parameter***

The input parameters of the numerical models (CCT, true IOP, SSI) and the output parameters (bIOP and CorVis deformation parameters) were used to derive an algorithm that provides estimates of SSI based on values of CCT, CorVis parameters and bIOP (the biomechanically-corrected IOP taken as a close representation of true IOP). The CorVis parameters were used first to provide values of the stiffness parameter at highest concavity (SP-HC). SP-HC was developed in an earlier study, is currently provided as a CorVis output, and has been shown to be strongly correlated to the cornea’s overall stiffness:

where *Adj*AP1 is the pressure measured at first applanation previously quantified using hot wire anemometry (Roberts et al., 2017); Deflectionmax is the amplitude of corneal apex deflection at the highest concavity; and DeflectionA1 is the deflection amplitude of corneal apex at first applanation. The least squares method was then used to develop an algorithm to determine SSI as a function of the numerical modelling input and output parameters; CCT, bIOP and SP. The method adopted the objective function:

Where *RMS* is the root mean square of the error, *N* is number of data points, *i* is the counter, *SSI Algorithm*is the value obtained from the algorithm, *SSI Numerical*is the value set in the numerical models

***Clinical Data and Validation***

The SSI algorithm was assessed against clinical data obtained from 480 healthy participants enrolled at the Vincieye Clinic in Milan, Italy (Dataset 1, 253 patients) and Corneal Tomography and Biomechanics Study Group – Rio de Janeiro, Brazil (Dataset 2, 227 patients). Institutional review boards at the two institutions ruled that approval was not needed for this record review study. However, ethical approval for using the data in research had been secured at both institutions when the data was collected, anonymised and used in earlier studies (Vinciguerra et al., 2016;Ambrosio et al., 2017b), before which participants’ informed and written consent was secured before collecting the data. Nevertheless, the ethical standards set out in the 1964 Declaration of Helsinki, and revised in 2000, were observed. All patients were evaluated with a complete ophthalmic examination, including the Corvis ST and Pentacam (OCULUS Optikgeräte GmbH; Wetzlar, Germany). All patients were free of any ophthalmic disease, with a Belin/Ambrósio Enhanced Ectasia total deviation index (BAD-D) derived from the Pentacam of less than 1.6 standard deviations (SD) from normative values in both eyes. Patients with previous ocular surgery or disease, myopia less than -10D, concurrent or previous glaucoma or hypotonic therapies were excluded.

All Corvis ST exams were acquired by the same experienced technicians with good quality (QS) scores that enabled calculation of all CorVis dynamic corneal response parameters (DCRs). Moreover, a frame-by-frame analysis of the exams, was performed by an independent masked examiner, to ensure quality of each acquisition. Only one eye per patient was randomly included in the analysis to avoid the bias of the relationship between bilateral eyes that could influence the analysis result. Any CorVis readings with visible rotational misalignment in the corneal profile were excluded from the analysis.

The clinical data were used to validate the SSI algorithm via testing the hypothesis that SSI would not be correlated with corneal thickness or IOP but be dependent on age (because of age’s correlation with material stiffness (Elsheikh et al., 2010b)).

***Ex-vivo Data and Validation***

As another form of validation, the correlation between SSI and age that has been established in the two clinical datasets was compared to what had been found in an earlier study involving inflation tests on ex-vivo human corneas (Girard et al., 2009;Elsheikh et al., 2010b). The study, which involved 57 corneas tested under inflation conditions with a posterior pressure simulating IOP, resulted in a stress-strain relationship of the form:

(Eq.4)

Where σ = stress, ε = strain, A = 1.26 × 10-3 and B = 0.0013 age2 + 0.013 age + 99. Both parameters of A and B are dimensionless. Differentiating Equation 1 with respect to the strain leads to:

(Eq.5)

where Et = tangent modulus. At the specific case with age = 50 years (at which SSI = 1.0), B = 102.9. Since the ratio between Et at any age and Et at age = 50 equals the ratio between SSI at this age and SSI at age 50 years, which is 1.0, therefore SSI at any age x can be determined from:

(Eq.6)

This value of SSI, based on ex-vivo results and given in terms of age, has been compared to the values of SSI obtained from analysis of the in vivo results, obtained from the two clinical datasets.

**Statistical Analysis**

Statistical analyses were carried out using IBM SPSS Statistics 24. Data were expressed as mean, standard deviation and range. Pearson correlation analysis was performed to study the relationships of corneal thickness (CCT), age and IOP with the SSI parameter. In this analysis, p values smaller than 0.05 were considered to be indicative of statistical significance.

**Results**

***Stress-Strain Index (SSI) algorithm***

The least squares method was used to develop an algorithm that can estimate the value of the SSI parameter based on the numerical modelling input and output parameters CCT, bIOP and SP-HC. The method resulted in a minimum RMS error of ±3% when the algorithm took the form:

SSI = *f* (a1 + a2 C1 + a3 C2 + a4 C12 + a5 C1 C2 + a6 C22 + a7 C13 + a8 C12 C2 + a9 C1 C22 + C23 + ln(SP-HC))

(Eq.7)

where C1 = CCT/545, C2 = bIOP/20, ln(SP-HC) the natural logarithm of the stiffness parameter at highest concavity, and a1 to a9 constants determined by fitting Equation 7 to the numerical input and output values, Table 1.

***Clinical Validation***

Dataset 1 (Milan)

Participants included in Dataset 1 had a mean age of 43.3±16.6 (8-87) years, CCT of 539.3±33.2 (454-629) microns and bIOP of 14.3±2.6 (7.7-29.3) mmHg. Analysis of CCT, bIOP, age and SSI values confirmed the hypothesis that SSI was not dependant on CCT (p= 0.792) or IOP (p= 0.745) but significantly correlated with age (P< 0.01), Figure 5. Statistical analysis was performed using Pearson correlation for bIOP and CCT as the data were normally distributed and with Spearman's rho correlation for age where the data were not normally distributed.

Dataset 2 (Rio)

In Dataset 2, participants had a mean age of 39.9±16.7 (7-81) years, CCT of 543.8±29.4 (454-621) microns and bIOP of 14.5±2.3 (9.8-24.3) mmHg. Similar to Dataset 1, the analysis showed that SSI was not dependant on CCT (p= 0.599) or bIOP (p= 0.281), but was significantly correlated with age (P< 0.01), Figure 6. Statistical analysis was performed using Pearson correlation with bIOP and CCT and Spearman's rho correlation with age for the reasons described above.

Combined Datasets

In order to increase the statistical power of results, the analysis was repeated while combining the two datasets. In this analysis, participants had a mean age of 40.6±17.1 (7-87) years, CCT of 541.5±32.43 (454-629) microns and bIOP of 14.7±2.4 (7.7-29.3) mmHg. Similar to the analysis conducted above, statistical comparisons showed that SSI was not dependant on CCT (p= 0.999) or bIOP (p= 0.480), but was significantly correlated with age (p< 0.01). The analysis was performed using Pearson correlation with bIOP and CCT and Spearman's rho correlation.

***Validation against ex-vivo inflation test results***

The relationship between SSI and age plotted in Figures 5c and 6c for Datasets 1 and 2, respectively, is re-plotted in Figure 7 and compared with the relationship based on ex-vivo inflation test results (Elsheikh et al., 2007). The comparison shows close correlation between the two relationships with the differences being 0.09±0.20 (p< 0.01) and 0.10±0.21 (p< 0.01) for Datasets 1 and 2, respectively. Statistical analysis was performed using Spearman's rho correlation as the data were not normally distributed.

**Discussion**

This paper attempts to address a long-standing challenge related to the in-vivo measurement of corneal biomechanics, and in doing so it attempts to overcome two major obstacles. First, the nonlinear nature of the tissue behaviour makes it necessary to determine the whole stress-strain behaviour, rather than a tangent modulus value which would be valid only at a particular level of stress or strain. This obstacle was overcome through an observation that stress-strain relationships obtained earlier for ex-vivo ocular tissue had similar trends that saw almost proportional decreases in strain with increases in tissue age. By taking the behaviour of corneal tissue at age 50 years as the benchmark, at which the new SSI parameter was assumed equal to 1.0, other stress-strain relationships for stiffer or softer material could be derived by multiplying the strain values by the relevant value of the SSI parameter.

The second challenge stems from the effect of IOP and corneal thickness on corneal deformation under the action of internal or external mechanical actions. However, while the effect of corneal thickness on overall behaviour is large, it can be estimated and removed as the thickness and its effect can be measured and excluded accurately. On the other hand, IOP presents a more difficult challenge since IOP measurement methods – through tonometry – are affected by corneal stiffness, creating a challenging dilemma with the stiffness affecting IOP measurement and IOP affecting corneal mechanical behaviour, which is used to estimate the stiffness. In this study, this challenge was addressed through consideration of a Corvis parameter – the stiffness parameter, SP-HC – which is more strongly correlated with corneal stiffness than IOP. For brevity, SSI is intended to be independent of IOP and corneal geometry and is needed to estimate the material stiffness, hence it is not the same as Stiffness Parameter (SP).

The new SSI algorithm was generated based on predictions of corneal behaviour using finite element (FE) numerical modelling simulating the effects of IOP and Corvis ST air puff. The algorithm was then validated through assessment of its correlation with IOP, CCT and age in two large clinical datasets. As expected, SSI was found to be independent of both IOP (p= 0.745 in Dataset 1, p= 0.281 in Dataset 2) and CCT (p= 0.792 in Dataset 1, p= 0.599 in Dataset 2), while being correlated with age (p< 0.01 in Dataset 1, p< 0.01 in Dataset 2), which, in turn, was found earlier – in an experimental study on ex-vivo human eyes) to be strongly associated with material stiffness (Elsheikh et al., 2007).

Another validation exercise was conducted by comparing the relationship between SSI and age established in the two datasets against the results of the earlier ex-vivo study (Elsheikh et al., 2010b). The comparisons showed there were no significant differences between the relationships (p< 0.01 in both Datasets 1 and 2).

The introduction of the SSI algorithms in clinical practice could enable customisation of the diagnosis and management of ocular diseases and allow optimisation of clinical procedures that either interact or interfere mechanically with the eye. With successful validation, SSI could help in identifying eyes with keratoconus, possibly increasing the sensitivity and specificity of indexes such as the Corvis Combined Biomechanical Index (CBI) (Vinciguerra et al., 2016) or the Tomography and Biomechanical Index (TBI) (Ambrosio et al., 2017b). Moreover, it could help in the detection of patients with higher risk or susceptibility for ectasia development or progression after refractive surgery and could aid in surgery planning. (Ambrosio et al., 2017a)

Glaucoma management could also benefit from the accurate measurement of corneal biomechanics (Kaushik et al., 2012). Among the factors that influence the accuracy of IOP measurement is the corneal tissue’s mechanical stiffness, and therefore quantifying the stiffness using the SSI algorithm could lead to improvements in IOP measurement and possibly better glaucoma management outcomes (Liu and Roberts, 2005).

There have been previous attempts to measure corneal mechanical properties in vivo. These included the Corneal Hysteresis (CH) and Corneal Resistance Factor (CRF) produced by the Ocular Response Analyzer (ORA) (Luce, 2005), and the Stiffness Parameter (SP) (Roberts et al., 2017) by the CorVis. These parameters were correlated with the diagnosis of keratoconus (Vinciguerra et al., 2016a) and showed significant increases after collagen cross-linking (CXL) (Bak-Nielsen et al., 2014) but could not provide measures of material behaviour that were separate from the effects of geometry and IOP. Another attempt is the elastic modulus provided by Brillouin microscopy (Scarcelli et al., 2013), which, while related to the cornea’s material stiffness, is not compatible with the nonlinear stress-strain behaviour that means the tissue does not have a unique modulus, but has a tangent modulus, which increases gradually with stress or applied pressure.

The SSI algorithm developed in this study is only suitable for corneas with normal topography. Corneas with keratoconus or ectasia, in which the geometry does not match the numerical models used in this work, will be treated separately in a future publication. Earlier work demonstrated the importance of including the ciliary muscles in simulations of corneal mechanical response to both IOP and external air pressure, but not the iris or the lens (Whitford, 2016). Earlier studies also confirmed the much lower stiffness of the retina relative to the ocular outer tunic (cornea and sclera) (Chen et al., 2010) and for this reason, it was not included in the numerical models.

In conclusion, we introduced in this study a new method for estimating the material behavior of healthy corneal tissue that can aid in optimisation of procedures that interact or interfere mechanically with the cornea..

**Author Contributions Statement**

KC conducted the research; AEliasy and RV drafted the manuscript, interpreted data, performed statistical analysis; AEliasy validated the findings and assisted in supervision; RV provided clinical data; AA, PV, RA, CR performed interpretation of data and provided clinical data. AElsheikh developed the concept, design the project and supervised the entire research and secured funding, all authors reviewed the manuscript and provided final approval

**Conflict of Interest**

RA, PV, RV, CR and A Elsheikh are consultants for OCULUS Optikgeräte GmbH. None of the remaining authors have financial disclosures. Prof. Elsheikh has received research funding from OCULUS Optikgeräte GmbH.

References

Ambrosio, R., Jr., Alonso, R.S., Luz, A., and Coca Velarde, L.G. (2006). Corneal-thickness spatial profile and corneal-volume distribution: tomographic indices to detect keratoconus. *J Cataract Refract Surg* 32**,** 1851-1859.

Ambrosio, R., Jr., Correia, F.F., Lopes, B., Salomao, M.Q., Luz, A., Dawson, D.G., Elsheikh, A., Vinciguerra, R., Vinciguerra, P., and Roberts, C.J. (2017a). Corneal Biomechanics in Ectatic Diseases: Refractive Surgery Implications. *Open Ophthalmol J* 11**,** 176-193.

Ambrosio, R., Jr., Lopes, B.T., Faria-Correia, F., Salomao, M.Q., Buhren, J., Roberts, C.J., Elsheikh, A., Vinciguerra, R., and Vinciguerra, P. (2017b). Integration of Scheimpflug-Based Corneal Tomography and Biomechanical Assessments for Enhancing Ectasia Detection. *J Refract Surg* 33**,** 434-443.

Avitabile, T., Marano, F., Uva, M.G., and Reibaldi, A. (1997). Evaluation of central and peripheral corneal thickness with ultrasound biomicroscopy in normal and keratoconic eyes. *Cornea* 16**,** 639-644.

Bak-Nielsen, S., Pedersen, I.B., Ivarsen, A., and Hjortdal, J. (2014). Dynamic Scheimpflug-based assessment of keratoconus and the effects of corneal cross-linking. *Journal of refractive surgery* 30**,** 408-414.

Belin, M.W., and Khachikian, S.S. (2006). New devices and clinical implications for measuring corneal thickness. *Clinical & experimental ophthalmology* 34**,** 729-731.

Chen, K.-J., Joda, A., Vinciguerra, R., Eliasy, A., Sefat, S.M.M., Kook, D., Geraghty, B., Roberts, C.J., and Elsheikh, A. (2018). Clinical evaluation of a new correction algorithm for dynamic Scheimpflug analyzer tonometry before and after laser in situ keratomileusis and small-incision lenticule extraction. *Journal of Cataract & Refractive Surgery*.

Chen, K., Rowley, A.P., and Weiland, J.D. (2010). Elastic properties of porcine ocular posterior soft tissues. *Journal of Biomedical Materials Research Part A: An Official Journal of The Society for Biomaterials, The Japanese Society for Biomaterials, and The Australian Society for Biomaterials and the Korean Society for Biomaterials* 93**,** 634-645.

Dubbelman, M., Weeber, H.A., Van Der Heijde, R.G., and Völker‐Dieben, H.J. (2002). Radius and asphericity of the posterior corneal surface determined by corrected Scheimpflug photography. *Acta Ophthalmologica Scandinavica* 80**,** 379-383.

Eliasy, A., Chen, K.-J., Vinciguerra, R., Maklad, O., Vinciguerra, P., Ambrósio, R., Roberts, C.J., and Elsheikh, A. (2018). Ex-vivo experimental validation of biomechanically-corrected intraocular pressure measurements on human eyes using the CorVis ST. *Experimental eye research* 175**,** 98-102.

Elsheikh, A., Alhasso, D., Kotecha, A., and Garway-Heath, D. (2009). Assessment of the ocular response analyzer as a tool for intraocular pressure measurement. *J Biomech Eng* 131**,** 081010.

Elsheikh, A., Geraghty, B., Alhasso, D., Knappett, J., Campanelli, M., and Rama, P. (2010a). Regional variation in the biomechanical properties of the human sclera. *Experimental eye research* 90**,** 624-633.

Elsheikh, A., Geraghty, B., Rama, P., Campanelli, M., and Meek, K.M. (2010b). Characterization of age-related variation in corneal biomechanical properties. *J R Soc Interface* 7**,** 1475-1485.

Elsheikh, A., Joda, A., Vinciguerra, R., Vinciguerra, P., Kook, D., Sefat, S., and Bao, F. (2015). Clinical Evaluation of Correction Algorithm for Corvis ST Tonometry. *Investigative Ophthalmology & Visual Science* 56**,** 101-101.

Elsheikh, A., Wang, D., Brown, M., Rama, P., Campanelli, M., and Pye, D. (2007). Assessment of corneal biomechanical properties and their variation with age. *Current eye research* 32**,** 11-19.

Elsheikh, A., Whitford, C., Hamarashid, R., Kassem, W., Joda, A., and Buchler, P. (2013). Stress free configuration of the human eye. *Med Eng Phys* 35**,** 211-216.

Ethier, C.R., Johnson, M., and Ruberti, J. (2004). Ocular biomechanics and biotransport. *Annu. Rev. Biomed. Eng.* 6**,** 249-273.

Geraghty, B., Jones, S.W., Rama, P., Akhtar, R., and Elsheikh, A. (2012). Age-related variations in the biomechanical properties of human sclera. *Journal of the mechanical behavior of biomedical materials* 16**,** 181-191.

Geraghty, B., Whitford, C., Boote, C., Akhtar, R., and Elsheikh, A. (2015). "Age-related variation in the biomechanical and structural properties of the corneo-scleral tunic," in *Mechanical Properties of Aging Soft Tissues*. Springer), 207-235.

Gilani, F., Cortese, M., Ambrósio Jr, R.R., Lopes, B., Ramos, I., Harvey, E.M., and Belin, M.W. (2013). Comprehensive anterior segment normal values generated by rotating Scheimpflug tomography. *Journal of Cataract & Refractive Surgery* 39**,** 1707-1712.

Girard, M.J., Suh, J.-K.F., Bottlang, M., Burgoyne, C.F., and Downs, J.C. (2009). Scleral biomechanics in the aging monkey eye. *Investigative ophthalmology & visual science* 50**,** 5226-5237.

Goldich, Y., Marcovich, A.L., Barkana, Y., Mandel, Y., Hirsh, A., Morad, Y., Avni, I., and Zadok, D. (2012). Clinical and corneal biomechanical changes after collagen cross-linking with riboflavin and UV irradiation in patients with progressive keratoconus: results after 2 years of follow-up. *Cornea* 31**,** 609-614.

Joda, A.A., Shervin, M.M., Kook, D., and Elsheikh, A. (2016). Development and validation of a correction equation for Corvis tonometry. *Comput Methods Biomech Biomed Engin* 19**,** 943-953.

Kaushik, S., Pandav, S.S., Banger, A., Aggarwal, K., and Gupta, A. (2012). Relationship between corneal biomechanical properties, central corneal thickness, and intraocular pressure across the spectrum of glaucoma. *American journal of ophthalmology* 153**,** 840-849. e842.

Liu, J., and Roberts, C.J. (2005). Influence of corneal biomechanical properties on intraocular pressure measurement: quantitative analysis. *Journal of Cataract & Refractive Surgery* 31**,** 146-155.

Luce, D.A. (2005). Determining in vivo biomechanical properties of the cornea with an ocular response analyzer. *Journal of Cataract & Refractive Surgery* 31**,** 156-162.

Nemeth, G., Hassan, Z., Csutak, A., Szalai, E., Berta, A., and Modis, L., Jr. (2013). Repeatability of ocular biomechanical data measurements with a Scheimpflug-based noncontact device on normal corneas. *J Refract Surg* 29**,** 558-563.

Ortiz, D., Piñero, D., Shabayek, M.H., Arnalich-Montiel, F., and Alió, J.L. (2007). Corneal biomechanical properties in normal, post-laser in situ keratomileusis, and keratoconic eyes. *Journal of Cataract & Refractive Surgery* 33**,** 1371-1375.

Pepose, J.S., Feigenbaum, S.K., Qazi, M.A., Sanderson, J.P., and Roberts, C.J. (2007). Changes in corneal biomechanics and intraocular pressure following LASIK using static, dynamic, and noncontact tonometry. *American journal of ophthalmology* 143**,** 39-47. e31.

Roberts, C. (2002). Biomechanics of the cornea and wavefront-guided laser refractive surgery. *Journal of Refractive Surgery* 18**,** S589-S592.

Roberts, C.J., Mahmoud, A.M., Bons, J.P., Hossain, A., Elsheikh, A., Vinciguerra, R., Vinciguerra, P., and Ambrosio, R., Jr. (2017). Introduction of Two Novel Stiffness Parameters and Interpretation of Air Puff-Induced Biomechanical Deformation Parameters With a Dynamic Scheimpflug Analyzer. *J Refract Surg* 33**,** 266-273.

Scarcelli, G., Kling, S., Quijano, E., Pineda, R., Marcos, S., and Yun, S.H. (2013). Brillouin microscopy of collagen crosslinking: noncontact depth-dependent analysis of corneal elastic modulus. *Investigative ophthalmology & visual science* 54**,** 1418-1425.

Villamarin, A., Roy, S., Hasballa, R., Vardoulis, O., Reymond, P., and Stergiopulos, N. (2012). 3D simulation of the aqueous flow in the human eye. *Medical engineering & physics* 34**,** 1462-1470.

Vinciguerra, R., Elsheikh, A., Roberts, C.J., Ambrósio, J.R., Sung Yong Kang, D., Lopes, B., Morenghi, E., Azzolini, C., and Vinciguerra, P. (2016). Detection of Keratoconus with a new Corvis ST Biomechanical Index. *J Refract Surg* 32**,** 803-810.

Whitford, C. (2016). *Biomechanical Properties of the Ocular Globe Based on Ex Vivo Testing and Multiscale Numerical Modelling.* PhD, University of Liverpool.

Whitford, C., Studer, H., Boote, C., Meek, K.M., and Elsheikh, A. (2015). Biomechanical model of the human cornea: considering shear stiffness and regional variation of collagen anisotropy and density. *Journal of the mechanical behavior of biomedical materials* 42**,** 76-87.

Table 1 Values of constants a1 to a9 used in Equation 7

|  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CCT(**µm**) |  |  |  |  |  |  |  |  |  |  |
| ≧ 400 | 0.3 | -3.094 | 5.249 | 8.982 | 0.248 | -8.423 | -2.416 | -0.443 | 1.704 | 2.198 |
| 0.5 | -7.731 | 22.224 | 7.699 | -17.455 | -8.806 | -1.515 | 5.361 | 2.852 | 1.471 |
| 0.7 | 0.440 | 0.387 | 4.723 | 2.974 | -5.498 | -0.403 | -1.200 | 2.386 | 0.404 |
| 0.8 | 4.509 | -10.507 | 3.013 | 12.998 | -3.028 | 0.017 | -4.315 | 1.583 | 0.002 |
| 0.9 | 7.603 | -17.995 | 0.764 | 18.971 | 0.888 | 0.297 | -5.826 | -0.114 | -0.259 |
| 1.0 | 8.047 | -18.217 | -0.500 | 18.236 | 3.236 | 0.395 | -5.235 | -1.242 | -0.336 |
| 1.5 | -8.355 | 30.668 | 1.754 | -30.649 | 0.651 | -0.519 | 11.572 | -1.163 | 0.653 |
| 2.0 | -3.101 | 16.284 | -0.219 | -18.494 | 4.480 | -0.208 | 9.073 | -3.482 | 0.508 |
| 2.5 | 4.677 | -9.969 | 3.607 | 10.742 | -1.410 | -1.504 | -1.413 | -1.463 | 1.804 |
| 3.0 | 6.842 | -16.245 | 3.244 | 17.519 | -4.064 | 0.222 | -3.391 | 1.251 | 0.092 |
| ≦ 400 | 0.3 | -3.094 | 5.249 | 8.982 | 0.248 | -8.423 | -2.416 | -0.443 | 1.704 | 2.198 |
| 0.5 | -7.731 | 22.224 | 7.699 | -17.455 | -8.806 | -1.515 | 5.361 | 2.852 | 1.471 |
| 0.7 | 0.440 | 0.387 | 4.723 | 2.974 | -5.498 | -0.403 | -1.200 | 2.386 | 0.404 |
| 0.8 | 4.509 | -10.507 | 3.013 | 12.998 | -3.028 | 0.017 | -4.315 | 1.583 | 0.002 |
| 0.9 | 7.603 | -17.995 | 0.764 | 18.971 | 0.888 | 0.297 | -5.826 | -0.114 | -0.259 |
| 1.0 | 8.047 | -18.217 | -0.500 | 18.236 | 3.236 | 0.395 | -5.235 | -1.242 | -0.336 |
| 2.0 | -3.101 | 16.284 | -0.219 | -18.494 | 4.480 | -0.208 | 9.073 | -3.482 | 0.508 |
| 2.5 | 4.677 | -9.969 | 3.607 | 10.742 | -1.410 | -1.504 | -1.413 | -1.463 | 1.804 |
| 3.0 | 6.842 | -16.245 | 3.244 | 17.519 | -4.064 | 0.222 | -3.391 | 1.251 | 0.092 |

CCT: Central Corneal Thickness; SSI: Stress-Strain Index;

Figure 1 Flowchart is demonstrating the process behind the analysis of built-in-house mesh generator software.

Figure 2 (a) A typical finite element model showing the boundary conditions applied at the equator and corneal apex and the four model regions, each with its own material behaviour. (b) Apex deformation of the numerical model during application of air-puff.

Figure 3 Spatial distribution (A) and temporal variation (B) of air pressure applied by the CorVis ST on the cornea(Joda et al., 2016)

Figure 4 Material biomechanical behaviour where (a) stress-strain curves intersect, or (b) stress-strain curves follow similar patterns. The almost linear variation of Et and SSI with applied pressure or stress, which corresponds to the behaviour patterns in (b) is depicted in (c).

Figure 5 Assessment of the correlation in Dataset 1 between SSI and each of (a) bIOP, (b) CCT and (c) age

Figure 6 Assessment of the correlation in Dataset 2 between SSI and each of (a) bIOP, (b) CCT and (c) age

Figure 7 Relationship between SSI and age based on in-vivo clinical data (black dots and a trend black line) and ex-vivo inflation test results (red dots) for (a) Milan dataset and (b) Rio dataset