# DEVELOPMENT OF A Beam Imaging System FOR THE European Spallation Source Tuning Dump

M. G. Ibison†, C. P. Welsch, Cockcroft Institute and University of Liverpool, UK

E. Adli, H. Gjersdal, G. Christoforo, University of Oslo, Norway** 
T. Shea, C. Thomas, ESS, Lund, Sweden

D. Naeem, STFC, Daresbury, UK

Abstract

To provide essential beam diagnostics for the spallation target and for the tuning dump of the European Spallation Source (ESS) with its high-power, low-emittance proton beam, optical imaging systems have been developed allowing remote viewing of the beam profiles, using scintillation light from coatings on the proton beam window (PBW) at the accelerator exit and the target entry window (TW), and insertable interceptive screens at the tuning dump (TD).

In this paper, we present the techniques used and the main parameters in the design of the TD system, including the optical configurations modelled in the software Zemax OpticStudio (ZOS). We also present the design principles for reflective optics which can transmit high-quality images, showing the performance of the imaging systems as predicted by ZOS and by prototyping. We take account of the requirements for initial alignment and ongoing maintenance of the optical system.

Studies of the radiation and thermal environments, which impact on sensitive optical components such as cameras and mirrors, are also described, as are mechanical considerations for the vacuum vessels and screen actuators. Finally, comments are made on the applicability of similar optics to diagnostic systems at other neutron sources and accelerators.

Keywords		Tuning dump, ESS, optics, diagnostics, radiation, modelling

TUNING DUMP optical diagnosticS: system requirements

The fundamental requirements for the TD diagnostics are summarised in Table 1. Beam conditions and the physical environment are specific to the dump line.

  - Tuning Dump Optical System Parameters

The Tuning Dump (TD) is designed to receive the ESS beam during initial commissioning and LINAC tune-up, whenever the beam is to be studied without sending it to the target. The dump can safely handle short proton pulses, and also full pulses but at reduced rate.

It was originally envisaged that a single imaging system located at the termination of the dump beam-line should be capable of viewing the beam incident on the dump face, in a similar way to beam-on-target imaging, as well as its profile immediately before its exit from the beam-pipe.

Table 1 TD system requirements - main parameters

| Parameter | Value | Origin |
| --- | --- | --- |
| Field of View | Max. possible | Full beam-pipe diameter = 250 mm |
| Limiting Apertures | 200 mm 100 mm | Vessel viewports Camera penetration* |
| Beam Size (nominal) | 1.6 cm (rms) | Beam dynamics simulation [1] |
| Resolution | <1 mm | Beam profile |
| Max Average Power | 12.5 kW | ‘Slow Tuning Mode’ [1] |
| Min Power | 0.1 kW | ‘Probe Beam Mode’ [1] |
| Min Energy | 200 MeV |  |
|  |  | *see Figure 14 |

  - Modified Tuning Dump System Requirements

It is no longer planned that the dump face will receive the luminescent coating which could have generated an image. The requirement for the TD system has therefore been simplified to providing imaging of the beam transverse profile at two or more locations upstream of the dump itself. These will be separated longitudinally to retain the capability for beam divergence assessment, which has been calculated as 10-3 under standard conditions; with a proposed separation of 1.7 m, the implied change of 1.7 mm in beam size will be measureable by a system achieving the required resolution. Positioning of the imaging stations in the final design is described at 5.1 and illustrated in Figure 12 below. The system is also required to be able to image a beam occupying any part of the full physical beam-pipe aperture at its exit.

optics DESIGN approach

Because of the hostile radiation environment in the TD tunnel area, a system based on all-reflective optics has been selected, rather than lenses or optical fibres; the latter are known to suffer rapid damage unless specially radiation-hard fused-silica fibre-bundles are specified as in [2], which give a design life of ≤ 108 R (~ 106 Gy). Unfortunately, these special fibres are no longer listed by the manufacturer as available. In any event, some degradation of components close to the proton beam is to be expected over time.

The leading optical system design software tool Zemax OpticStudio [3] has been used to perform all design work, with theoretical values as initial parameters. While not essential, the use of Zemax has assisted with integration of the optics with beam-line CAD models. Pure ray-optics modelling was considered to be sufficiently accurate, as diffractive effects may be neglected for the TD optical system.

  - Tuning Dump Optics Design

Initial concepts explored for TD imaging included a 3-curved-mirror system based on aberration compensation principles [4], resembling the optics previously designed for PBW/TW imaging [5][6][7], and also an optical fibre solution with cameras located remotely. The multi-mirror option was considered unnecessarily complex for the minor gain in field-of-view, while fibres would be too susceptible to radiation damage and have inferior optical characteristics, in addition to cost and availability issues.

The final design selected has a simple ‘periscope’ configuration comprising 2 plane mirrors, which combines acceptable image quality with flexible camera positioning.

To meet the original requirement, which envisaged the imaging of the beam as it entered the dump face, an early design was devised with a composite screen/mirror system to image the beam on the screen front surface as it passed through, while simultaneously collecting a reflected image of the beam on dump by using the rear surface of the screen as a mirror. With the relaxation of the dump imaging requirement, it has become possible to simplify the system to a plain screen, eliminating the need for rear mirroring.

The optical components comprising the model therefore include:

- the object (screen intercepting proton beam)

- the viewport in the vacuum vessel

- 1st 45° mirror (outside the viewport)

- 2nd 45° mirror (on ray-path from 1st mirror)

- imaging lens and camera

Figure 1 Screen in centre of imaging vessel - plan view. Proton beam direction is indicated by the arrow; the screen orientation shown is at 90° to the beam.

  - Photon Sources for Imaging

The primary light source for TD imaging will be a ‘Chromox’ ceramic screen as shown in Figure 1, although an alternative having a thermally-sprayed luminescent coating, most probably Cr-doped alumina on an aluminium backing, will also be fitted. Screen materials are excited into photon emission by the energy of incident protons. Studies are ongoing to identify materials which could improve on the standard widely-used and commercially-available ‘Chromox’ [7]; however, these must give adequate photon yield per proton, emission spectrum, lifetime & linearity, while maintaining their properties after the heat exposure of the spraying process. Although the continuing research into photon sources will be published separately, a brief analysis of the predicted intensity is presented here for completeness.

For a thin 1mm Chromox screen, the expected light yield is 5 x 104 photons per proton, given an average loss of 1 MeV/mm, principally as emission lines at 691 and 694 nm with lifetimes of 6.7 and 3.4 ms. A minimum-power 6.25 mA ‘probe’ beam (see Table 1) sent to the dump would therefore produce some 1016 photons per 5 s pulse. For the nominal beam-size of 1.6 cm, 1.7 x 1013 photons/mm2/s will reach the 2048 x 2048 camera sensor, assuming that the optics are set to image the full screen width of 32 cm.

Taking the specified camera sensitivity of 14.2 e- per 5.5 m pixel and 50% quantum efficiency, an irradiance of at least 106 photons/mm2 is required on the sensor for imaging. The screen is therefore estimated to provide this level with exposure times of <1 s, and will continue to give adequate light even with the efficiency loss experienced after prolonged irradiation.

Figure 2 FLUKA model of the dump line tunnel (elevation shown from RHS); beam from L to R. Dimensions in m.

studies of PHYSICAL ENVIRONMENT

The physical environment in which the optical systems must function has an effect not only on their performance, in terms of image quality and stability, but also on the useful lifetime of their components. This applies to the effects of both heating by deposition of energy from the beam and also damage at the atomic level from particle collisions.

  - Tuning Dump – Radiation Environment

The dump line with its vacuum vessels, the dump structure itself, and the enclosing shielded tunnel, have been modelled using the Monte-Carlo radiation transport code FLUKA [9]. By applying the expected beam parameters at the dump and the estimated utilisation, based on the ESS operating schedule, to the outputs of the model, the total radiation dose per year has been calculated, at locations of interest in the vicinity of the imaging vessel. The objective was to find positions for the cameras which while providing the required field of view would ensure an adequate useful lifetime, before radiation damage to the sensor compromised the image quality unacceptably.

  - Camera Radiosensitivity

Studies have been made both of dose and of neutron or proton fluence, as related to observable effects expected in image quality or system performance, including:

- permanent damage to sensor pixels

- ‘upsets’ to camera electronics, recoverable after reset

- irrecoverable damage causing imaging failure

Based on these sources [10], and the expected major annual shutdown schedule during which degraded optical elements may be replaced, an arbitrary target dose of < 20 Grays/year has been set for selection of the imaging camera locations, with any associated shielding. At a fraction of this dose, no detectable effects are expected.

A number of alternative, specifically ‘radiation-hard’ camera solutions were considered, such as the older ‘Vidicon’ tube type which are still available (e.g. Mirion), or CID sensor types (e.g. Thermo Scientific).

These have quoted operational dose limits of 2x106 and 3x104 Gy (lifetime), and 3x104 and 1x103 Gy/hr (during imaging), respectively. However, for optimum image quality and choice of supply, an adequately shielded standard scientific camera is preferred, specifically the ‘Manta G-419’ by Allied Vision; also listed in Table 5 below are details of the other main optical elements. As occasional single event upsets from radiation are to be expected, the camera control system must allow rapid reset by remote power-cycling.

Although no other critical component is as radiosensitive as the cameras, dose assessments have also informed the selection - or avoidance - of such susceptible materials as elastomer vacuum seals and actuator flexible drive belts.

  - Radiation Transport Modelling with FLUKA

The FLUKA code is a mature simulation environment for the Monte-Carlo modelling of radiation transport (charged particles, neutrons & photons) through matter. A model of the ESS tuning dump, the TD beam-line and its shielded tunnel has been created, including all significant geometry and material compositions, as shown in Figure 2. A simulation is run by propagating particles from the start of the beam-line towards the dump and collecting data on the history of each primary particle as it interacts, possibly producing secondary particles which are also tracked. As events in the regions of particular interest were rare, a very large number of primaries had to be followed; typically, this was 5x105 to 106 for the TD geometry, for statistically meaningful results to be obtained. Following FLUKA guidelines, five repeated independent runs were performed for each model, allowing an estimation of the associated variance.

The external viewer program ‘SimpleGeo’ has proved useful for preparing detailed 3-D geometry images and for the plotting of particle tracks, based on FLUKA data files; it also provides a CAD modelling interface.

Separate studies were carried out to evaluate doses or dose-rates for three different cases:

- Dose to camera while the ESS beam is on the dump rather than the target

- Dose from particle losses into the dump tunnel, during ‘User’ (beam-on-target) operation

- Dose-rates at the imaging station after beam-on-dump operation, due to residual proton-induced radioactivity

  - Camera Radiation Dose

For the TD system, the absorbed dose was recorded in those regions proposed for locating the imaging cameras. Dose is estimated by FLUKA per primary particle, the total number of which must be calculated from the projected beam current and the annual time for beam on dump, based on the equation:

Protons per year = 0.5 x tS x 3600 xIP e = 3.54 x 1019

where annual machine study time tS = 500h; time-on-dump fraction (estimated) = 0.5; beam current (mean) IP = 6.3A; e = electronic charge

It was also demonstrated that a large fraction of the dose received by an unshielded camera is due to particles scattered by the imaging screen itself; with the screen removed, only that radiation escaping from the dump entrance makes a contribution. Figure 3 shows the imaging vessel in the dump-line tunnel, with the proton beam propagating from L to R. Tracks of different types of scattered particle are indicated by colour. On the top picture, the screen is inserted into the beam; on the bottom it is removed. Each run is for equal numbers of primary protons.

As new candidate locations were explored, the FLUKA model was developed progressively by adding the required geometry, so that doses at multiple camera positions could be compared during the same run.

Figure 3   FLUKA particle tracking displayed in ‘SimpleGeo’, with screen inserted (top) and removed (bottom). Camera locations shown are unshielded. Beam is L. to R.

Table 2 Camera doses for different locations. Camera to R. of beam. Doses in Gy/year; errors at ±1.

| Location | Camera Dose | Position x | Position y | Position z | Screen (x,y,z) |
| --- | --- | --- | --- | --- | --- |
| Shielded Bunker on Tunnel Floor, 5cm concrete lid | 380 ± 90 | -24.2 | -37.5 | -524.2 | (0,0,-500) |
| Shielded Bunker, thicker cover plate, 7.5cm concrete | 650 ± 140 | -24.2 | -37.5 | -524.2 | (0,0,-500) |
| Shielded Trench excavated in Floor Slab, 7.5cm cover | 275 ± 350 | -24.2 | -67.5 | -524.2 | (0,0,-500) |
| High-Level†, otherwise layout as for Bunker case | 40 ± 70 | -24.2 | 105 | -524.2 | (0,0,-500) |
| High-Level†, further displaced from beam axis | 24 ± 22 | -35 | 105 | -535 | (0,0,-500) |
| High-Level†, down-looking onto 1st mirrors | 55 ± 30 | -45 | 120 | -545 | (0,0,-500) |
| High-Level, horizontal in wall ‘niches’ | 15 ± 9 (3 ± 5)* | -78.5 | 125 | -579 | (0,0,-500) |
| High-Level, set well back into walls | 12 ± 7 | -100 | 155 | -633 | (0,0,-533) |
| High-Level, in 1.5m hole | undetectable | -215 | 155 | -601.2 | (0,0,-561.2) |
| *Note: Doses with screen removed from beam; see Figure 3. †High-level camera unshielded, unless in walls. |  | Camera mid front location in cm. Origin of coordinates is centre of dump face: transverse x, vertical y, longitudinal z |  |  | Screen centre coordinates. |

In Table 2, results show that with the camera located close to or just under the beam-line, addition of shielding alone cannot sufficiently reduce the dose to the target value. One possible approach to the problem could have been to move the camera to a remote position with an acceptable dose-rate; however, allocating the long, clear light-path required would not have been feasible in the restricted space of the dump tunnel, and also the light collected from the screen would be much reduced. The combination of height and the optimal use of the existing shield wall thickness has provided a solution. Drilling restrictions limit the size of hole, primarily due to the wall

Figure 4  Camera location (final) as listed in Table 2. Plan view from above, in a horizontal plane through the camera. Dimensions in cm.

reinforcement, but the hole will still have a minimum diameter of 100mm; its depth will not exceed 1.5m. It should also be noted that due to changed imaging requirements (see 1.2 above), only a single camera will in fact be deployed, as listed in Table 2. Figure 4 illustrates the camera position in the final design.

  - Dose from Beam Losses

When the ESS is operating in ‘User’ mode, the beam from the LINAC is directed into the target line by dipole magnets in the first bend of the ‘dog-leg’ section. Some of the protons are lost from the beam in this region and will enter the dump tunnel, as shown in Figure 5, contributing to the total dose received by the imaging cameras. To model these losses, advantage was taken of earlier ESS modelling work to provide the input data files containing full parameter sets (position, direction vectors and energy) for a large set of particles. A customised code module ‘SOURCE’ was written to read the pre-processed particle data into the existing TD FLUKA model. From this, the dose per source proton at the camera location was obtained, and hence the annual dose derived using the following result:

Proton Losses to Dump Tunnel per year=(5300 ×3600) ×(2.52 ×10-3)×(0.002 ×0.01) / e = 6.00× 1018 protons

where average beam current (5MW full power beam on target) = 2.52mA, operating hours = 5300 per year, and fractional loss rate at the dipole = 0.002% (assumed).

Doses to the camera in its final selected location as shown in Figure 14 were undetectable in FLUKA simulations.

Figure 5  Beam losses from first dipole in the dog-leg section, which enter the dump tunnel. Beam from LINAC enters from L, beam to target leaves to R.

  - Decay Dose-Rates

A separate study was made of the dose-rates which would be expected in the vicinity of the imaging station, close to the TD, in the period immediately after beam shut-off, due to the decay of activation products in and around the dump.

The pattern of beam-on-dump periods taken from the proposed ESS Operations Schedule, after the initial start-up date, is shown in Figure 6.

Figure 6  Projected beam-on-dump irradiation profile for 1 full year, after the start of normal user operations. Particles delivered in each period are based on an average beam current Ip = 6.3A.

FLUKA accepts arbitrary particle irradiation times and rates as input, and provides dose-rate output at any selected decay times after beaming. Dose-rate 2-D profiles in the horizontal plane of the beam are plotted in Figure 7, for times of 1, 8, 24 and 72 hours after shut-down, at the end of one year’s operation.

Figure 7  Decay dose-rate plots in horizontal beam plane in the vicinity of the imaging vessel, after 1, 8, 24, & 72 hours’ cooling time. Distances in cm, dose-rate in Sv/h.

The intense spot in the centre-right of Figure 7 is due to the imaging screen; decay in radiation streaming from the activated dump itself (to R of the plots) is also evident.

An independent study was made by analytical calculation of the activity induced by a 4.5cm radius beam of 2GeV protons via Cu(p,xn) reactions [11] in the copper dump cylinder and the subsequent decay of the 21 most important nuclides produced, using the equation [12]:

SABq g-1=NAσA dφdt 1-e-tiln2t12e-tcln2t12

where NA is Avogadro’s Number,  is the cross-section for the reaction, A is the nucleon number for Cu (taken as the atomic weight for natural copper), d/dt is the particle fluence rate, ti is the irradiation time, tc is the cooling time, and t1/2 is the radioactive half-life. This equation gives the activity per gram SA at time tc after the end of the irradiation. The result is converted into the total activity in the volume of the dump, using the density of Cu.

The gamma dose-rate at 4m from this source on the beam axis (the approximate location of the imaging vessel), was then estimated. The dose-rate D’ is calculated for each significant gamma-ray from [13]:

D'=AnE4πμenρG

where A is the nuclide activity, n is the gamma-ray abundance, E is the gamma energy (in J), en/ is the mass energy-absorption coefficient (for air), and G is the ‘geometry factor’ for a cylindrical source, derived from the expression [14]:

1R02hh2+zln1+R02h2+z2-z-h2ln1+R02z-h22+2R0tan-1h2+zR0-tan-1z-h2R0

where R0 and h are the radius and height of the cylinder, and z is the distance from its centre, on axis.

The contributions are then summed for the total dose-rate.

Table 3 Summary of decay dose-rates from two estimation methods. All dose-rates are quoted in mSv/hr.

| Cooling Time (hours) | 1 | 72 |
| --- | --- | --- |
| Total Dose-Rate (analytical) | 51.3 | 20.6 |
| Dose-Rate (from FLUKA) | 10-100 | 1-10 |

The results shown in Table 3 are consistent, given the calculation uncertainties, especially as the analytical approach ignores self-shielding by the dump. At the shorter cooling time, FLUKA data will be enhanced by rapidly- decaying local radiation from the screen.

In summary, it has been shown that for the assessment of expected camera lifetime, the contribution to the dose from scattered radiation during beam-on-dump may be very effectively controlled by shielding provided by the tunnel wall even at shallow depth, and is negligible at 1.5 m (in Section 3.4); while the dose from beam losses (Section 3.5) may be neglected entirely. The decay dose-rate shortly after beaming is significant enough to limit manned access time to some minutes (Section 3.6), especially when the ‘cooling’ period is much less than 72 hours.

  - Screen Heating Studies

Figure 8  Model of passage of beam through screen, showing 'core' region for thermal analysis (1 width).

As the beam passes through the imaging screen as in Figure 8, it deposits some energy as heat, thereby raising the temperature [15]. In the final design of Figure 14 the beam is orthogonal to the screen, but the heat deposited per unit volume is still approximately the same.

In the beam-pipe vacuum, only radiative and conductive processes are available to dissipate the heat. Studies were carried out into the instantaneous heating effect of a single full ESS proton pulse, assuming no immediate heat removal. In Figure 9, the peak temperature reached in the various layers of a composite screen is plotted against beam size (at 1 assuming a Gaussian distribution) for a beam energy of 570 MeV, this representing the most severe heating regime expected [1]. For this purpose, beam ‘core’ regions (1) of the screen layers are defined, treated as if in thermal isolation from the outer regions.

Figure 9  Peak temperature in screen layers vs. beam size, for a single full pulse of 1.114x1015 protons, at Ep = 570 MeV.

Considering the melting points (MP) of candidate screen materials, these results indicate that a coated aluminium–alloy screen (MP ~ 580°C) is safe down to a beam size x ≥ 0.75 cm, while a Chromox screen (MP ~ 2000°C) could be used at x ≥ 0.4 cm.

In comparison, a nominal beam of x = 1.6 cm is predicted to heat the screen materials by a maximum of 70°C. It is recognised that the light emission from a heated screen is temperature-dependent, and this effect is considered in ongoing studies of screen material in proton beams; these are still in progress and will be reported at a later date.

system prototyping

  - Tuning Dump System

Figure 10 Prototyping with optical table for resolution tests, showing light path to camera via mirrors M1, M2 (looking down on path, camera seen in L. foreground).

At the University of Oslo (UiO) a prototyping facility has been constructed with two large aligned optical tables as shown in Figure 10; these may be used in tandem to model the long path lengths of the ESS optical systems. Prototyping assists in confirming the results of ZEMAX optical simulations, which cannot always reliably incorporate the effects on image quality of realistic errors in optical components. The effects of apertures and limits on fields of view can also be evaluated, and the proposed initial assembly, alignment and maintenance procedures assessed.

It should be noted that as seen in Figure 10, the vertical leg of the optical path in the Final Design has been translated into the horizontal, for convenience; but this has no influence on demonstrating the effectiveness of the imaging. The prototype has provided the basis for the realistic assessment of optical performance, as illustrated in the image of Figure 11 below showing that 1mm graduations are easily resolvable; the results in terms of optical parameters are summarised in Table 4.

Description of final Design

  - Optical Performance of the TD Design

Figure 11a   MTF Plot of full TD optical system.

Conventionally, the Modulation Transfer Function (MTF) is the standard measure of optical performance in imaging fine detail, which is plotted for the ZEMAX model of the TD system in Figure 11a. It is expressed as the image-to-object modulation intensity ratio of closely-spaced light/dark sinusoidal cycles, over a range of spatial frequencies. However, a more realistic prediction of image quality than that simulated by Zemax has been obtained experimentally at the prototyping facility, equipped with a lens, mirror and camera identical to those planned for installation, as shown in Figure 11b.

Figure 11b  Image quality assessed by viewing a millimetre resolution test scale at the prototyping facility, through the same camera and lens as those proposed, plus a sample of viewport window.

In view just R. of centre is a small sample representative of the proposed viewport window, through which the 1 mm graduations of a test scale at the screen position are shown to be well resolved. Figure 11c illustrates a standard optical system test target, with its simulated image as viewed through the proposed Tuning Dump optics. The image inversion is a natural property of the mirror/lens system. Although minor distortion is evident, and a degree of vignetting is evident at the extremities, the resolution fully meets the imaging requirement.

Figure 11c   USAF 1951 test target L., with its simulated image seen through the tuning dump optical system R.

Figure 12   Locations of imaging stations in the dump line. Tertiary imaging vessel is to be initially installed empty. Beam direction is from bottom L to top R.

The limiting aperture in the system is the 200mm diameter viewport, the mirrors being sized to accept the full field of view. The main optical properties of the chosen design are listed in Table 4; these are the values expressed in the Zemax model.

Table 4 Design parameters for the Tuning Dump system - values entered into the Zemax model

| Parameter | TD system |
| --- | --- |
| Focal Length (mm) – set by lens selected | 135 |
| f/# {proposed} | f/2.85 {f/2}† |
| Mirrors: Clear Diameter (mm) M1 M2 | 290 110 x 150 |
| Screen – Lens Distance (mm) | 3828* |
| *variable, depending on exact position of camera † depends on position & hole diameter (see Figure 14) |  |

  - Tuning Dump System Components

Initially, two identical systems will be installed, with provision for a third at a predefined upstream location, indicated in Figure 12; access to three profiles would enable more advanced diagnostics, including emittance measurements.

    - Mechanical System

Each imaging vessel will be special custom-made 5-way cross, the horizontal arms conforming to the DN-250 standard of the beam-pipe, while the vertical will be to DN-350 to accommodate screens wide enough to allow coverage of the full aperture, illustrated in Figure 13. With a beam-height of 500mm above the tunnel floor, there is space below the beam-line for a vessel sized to accept the one unused screen in its lower vertical section, while the upper part is sufficiently tall to take both screens when fully raised clear of the beam.

To facilitate maintenance reassembly, threaded ‘half-moon’ split-rings, rather than separate nuts (x36 for DN-350 size), will be used for clamping the top and the 45° viewport flanges of the vessel; this also allows the flanges to be rotatable.

Figure 13 Imaging vessel final design (axiometric view). Beam enters from lower R., viewport flange to front. (Inset: Actuator showing limit switches.)

A long-travel vertical linear actuator on the top flange, also seen in Figure 13, can move either of 2 screens into the beam. The actuator is a conventional edge-welded bellows design driven by a lead-screw. Special features include an in-line planetary gearbox to meet the torque requirement of 2 N-m from a stepper motor with a moderate current load, though a directly-coupled option is also under consideration; the motor is selected for ‘harsh-environment’ (radiation) service. The configuration also avoids the need for an elastomeric drive-belt, which could be susceptible to radiation-induced failure; the annual dose to the motor, which is at 2.1 m above floor level and therefore well away from the primary beam, has been assessed by FLUKA modelling as 50 Gy. The system has no encoders, motion being controlled by five limit switches which sense intermediate screen positions as well as ends-of-travel.

    - Optical System

Screens will be mounted at 90° to the beam direction and will be viewed directly through a large diameter fused-quartz viewport on the vessel’s 45° arm. The viewport window may be changed when needed due to transmission loss from radiation darkening; although quartz is radiation-resistant, FLUKA modelling predicts a significant dose of up to 5.7 kGy/year, mainly from proton scatter, as the window is at beam height and only 30 cm off-axis.

Images will be relayed vertically through pairs of 45° plane mirrors, to cameras mounted in 1.5m deep holes drilled in the tunnel walls, 1.5m above beam height, as shown in Figure 14. Remotely-controlled lenses on the cameras will focus the final images and also provide for adjusting the f-number. Remote filter-changers located just in front of the lens, or alternatively at the entrance to the shield-wall hole leading to the camera, will provide selectable attenuation for image intensities which could, under certain beam-focusing conditions, saturate the camera. It is possible to correct the depth-of-field error across the full screen width, as viewed at a 45° angle, by applying the Scheimpflug principle [16]. In the equation   tanψ= u'fsinθ,  with object angle ψ = 45°, lens focal length f = 0.1 m, and object distance u’ = 4 m, the predicted tilt θ required in the camera sensor plane is 1.43°.

Characteristics of the selected optical elements are listed in Table 5.

An interface with the Machine Protection System, to inhibit the beam in case an unsafe condition is detected, will be provided through logic in the image processing software [17].

    - Control Systems

Controllers for both actuators and cameras will be located in racks outside the dump tunnel area. Motion control will use the ESS standard EtherCAT protocol which is supported by a range of commercially-available modules; these will also provide the logic processing for the limit switch array.

Figure 14  Optical path from a screen inside the vessel, via plane mirrors to a camera located inside a hole drilled into the shield wall. Proton beam direction is indicated by the red arrow. (Inset: Plan view, with beam from bottom.)

Table 5 Selection of Principal Optical Components. Mirror surface quality is quoted in wavelengths.

| Component | Property |
| --- | --- |
| - Mirror M1 | UQG Optics 290mm diam x 6mm thick Aluminium surface coated on borosilicate; SiO2 protected Quality: 1  per 25mm. |
| Mirror M2 | UQG Optics 150 x 110mm x 10mm thick Silver coated on UV fused silica; SiO2 protected Quality: /4 across surface Hardness: 40-20 scratch-dig |
| Lens | Canon 135mm f/2L USM |
| Camera | Allied Vision Manta G-419 Sensor: CMOS 2048 x 2048 pixels |

Figure 14 was generated in the ‘non-sequential’ mode of Zemax, which has also been used to explore the options for providing illumination of the screen in the absence of a beam. For this, it is currently proposed to locate a radiation-hard light source between mirror M1 and the viewport.

CONCLUSION

An optical system has been designed to form images of the ESS proton beam profile at the end of the Tuning Dump line. The design has been developed and optimised to meet performance requirements under severe radiation environment constraints, and corroborated using the Zemax toolset. It has been shown that it is possible to adopt a very simple two-plane-mirror and lens system while maintaining acceptable image quality; the cost and complexity of improving the image using further mirrors is not considered justified. Although detailed tolerancing studies for the effects of thermal expansion, misalignments, and production uncertainties are not presented here, prototyping of the system with mirrors of the specification to be installed has demonstrated they are able to meet imaging requirements in terms of field-of-view, to observe the beam at any point on a fluorescent screen covering the full aperture, and to resolve details to better than 1 mm. In addition, depth-of-field correction has been made, despite a 45° viewing angle, with minimal distortion. These results have been supported by optical simulation studies.

Assessment of predicted radiation dose in the Tuning Dump line after irradiation has informed the location of imaging cameras for acceptable working lifetimes, the choice of materials for other key components, and the conditions to be expected during maintenance access. A detailed design for the imaging vacuum vessel and its mechanical elements has been developed to meet vacuum and other requirements.

Applicability to Other Neutron Sources

The design criteria have emphasised resilience and durability, giving assurance of longevity together with maintainability. As such, an imaging system of this type would be suitable for other high-power proton beamlines, unless a non-invasive diagnostic is required. Where there are pre-existing shielding structures in the beam-line design, it may be more practical to make use of them rather than to attempt to retro-fit additional shields specifically for the optical diagnostic system. However, this approach assumes that a sufficiently realistic model of the beam-line, and adequate beam parameter prediction data, is available for dose assessment in the relevant regions of space. The adoption of simple non-refractive optics leads to a robust system with reduced costs; plane mirrors, even at large apertures and high optical qualities, are relatively inexpensive and are also easier to mount and to align in situ than lenses or curved mirrors, which is important in restricted access, high-radiation environments. In similar situations, the approach described here may be preferred over the greater complexity and possible cost of systems such as that covered in [2].

The element of the system carrying the greatest risk is the screen itself, which is subject to possible beam damage, breakage or loss of emission; however, this is mitigated by provision of a running spare at each imaging station.

ACKNOWLEDGEMENTS

This work is supported by the In-Kind Agreement between ESS ERIC and the University of Oslo, Norway.

The author would also like to thank students Scott Doyle and Ellis Chaffe for their invaluable assistance in compiling the analytical dose-rate data presented in Table 3.

References

[1] Y. Lee, A. Olsson, M. Eshraqi et al., “Working Concept of 12.5 Kw Tuning Dump at ESS”, in Proc IPAC2017, Copenhagen, Denmark, paper THPVA065, pp 4591-4

[2] L Maxey, T Ally, A Brunson, F Garcia, T Shea et al., “A hybrid reflective/refractive/diffractive achromatic fiber-coupled radiation resistant imaging system for use in the Spallation Neutron Source (SNS)”, Proc. Hard X-Ray, Gamma-Ray, and Neutron Detector Physics XIII, Sept 2011, SPIE Vol. 8142, 81420N-1

[3]	ZEMAX LLC, Kirkland, WA 98033 USA, www.zemax.com/os/opticstudio

[4]	D Korsch, “Reflective Optics”, San Diego, USA: Academic Press, 1991, ISBN: 978-0-12-421170-4.

[5]	M Gohran, R Linander, T Grandsaert, T Shea, C Thomas, “Proton Beam Imaging Options for the ESS Target”, in Proc. IPAC 2014, Dresden, Germany, June 2014, paper THPME168, pp 3659-3661.

[6]	M Ibison, C. P. Welsch, E. Adli, H. Gjersdal, N. de la Cour, T. Shea, C. Thomas, “Optical System Design for The ESS Proton Beam and Target Diagnostics”, in Proc. IPAC 2016, Busan, Korea, paper MOPMR043, pp 347-349

[7]	E. Adli, H. Gjersdal, O. M. Røhne, C. Thomas, T. Shea. R. Andersson, M.G. Ibison, C.P. Welsch, S. Joshi, “The ESS Target Proton Beam Imaging System as In-Kind Contribution” in Proc IPAC2017, Copenhagen, Denmark, paper WEPVA066, pp 3422-4

[8]	M Donna, T Grandsaert, M Göhran, R Linander, T Shea et al., “Optical System for ESS Target Protection”, in Proc. IBIC 2014, Monterey, CA, USA, Sept 2014, paper TUPF27

[9] A. Ferrari et al., “FLUKA: a multi-particle transport code”, CERN-2005-10, INFN/TC_05/11, SLAC-R773, 2005.

[10] S. Hutchins, M. Facchini, E. Tsoulou, “Radiation Tests on Solid State Cameras for Instrumentation”, Proc. DIPAC 2005, Lyon, France

[11] R. Michel, M. Gloris, H.-J. Lange et al, “Nuclide production by proton-induced reactions on elements (6 ≤ Z ≤ 29) in the energy range from 800 to 2600 MeV”, Nuclear Instruments and Methods in Physics Research B 103 (1995) 183-222

[12] Bruce C. Brown, Nikolai V. Mokhov and Vitaly S. Pronskikh, “Activation Products from Copper and Steel Samples Exposed to Showers Produced by 8 GeV Protons Lost in the Fermilab Main Injector Collimation System”, Fermi National Accelerator Laboratory, Batavia, IL, USA

[13] “Dose Calculations: Absorbed Dose from a charged particle beam”, MIT Course in Nuclear Engineering, Principles of Radiation Interactions 22.55 (2004)

[14] Ana Neacsu, Mihail Contineanu, Traian Zaharescu, Iulia Contineanu, “Calculation of the Gamma Radiation Dose Produced by a Cylindrical Radioactive Source”, REV. CHIM.(Bucharest) 67 No. 9  2016

[15] M. Turner, B. Biskup, S. Burger, E. Gschwendtner, S. Mazzoni, A. Petrenko, “The two-screen measurement setup to indirectly measure proton beam self-modulation in AWAKE”, Nuclear Instruments and Methods in Physics Research A 854 (2017) 100–106

[16] T. Scheimpflug, “Improved Method and Apparatus for the Systematic Alteration or Distortion of Plane Pictures and Images by Means of Lenses and Mirrors for Photography and for other purposes”, UK Patent No. 1196, Jan 1904

[17] T Shea, L Coney, R Linander, A Jansson, E Pitcher et al., “Instrumentation & machine protection strategy for the ESS target station”, in Proc. ICANS XXI, Mito, Japan, Sept 2014